## Instructions
FFDM Medical Physicist QC Forms
This set of worksheets is designed to be used on a laptop or computer, directly entering the data as acquired.
The sheets are set up to be performed in the sequence specified in the manual, in order to reduce the number of images, and duplicate set-ups.
If tests are performed in the correct order, the data will be automatically inserted in the spreadsheet at appropriate locations.
The spreadsheets may be filled out by hand, but automatic calculations are not performed.
The chart numbers apply to the procedure number associated with that number; however, a number of charts incorporate data from more than one procedure.
Cells which need to be filled in while at the acquisition station are identified by being shaded like this:
Cells which are automatically calculated or filled in from a previous worksheet are identified by being shaded like this:
Cells which are filled in at the review workstation are identified by being shaded like this:
Handy navigation macros:
Ctrl-a | A-Raw Image Acquisition
Ctrl-l | Monitor Lum. Response 14
Ctrl-t | Thickness Tracking 3
Ctrl-i | CR Interplate Sensitivity 18
Harvesting Data.
Kodak CR Images
# Image Harvest: the 5.x version of the acquisition workstation software makes it very simple to harvest the images.
It's easiest if you DON'T approve your images after taking them (i.e. the border of the thumbnail versions remains green)
because, otherwise, you need to do more searching to find them instead of just querying, say, Lastname = "PHYSICS".
# Incorrect OrganDose: the same value is currently stored in OrganDose in every image. Andja from Carestream has been informed.
# Processed Images: only unprocessed images may be harvested directly from the CR Reader.
If you want to obtain processed versions of images, you need to get them from a point after they've been sent from the Acquisition Workstation (i.e. the PACS).
# Key Operator Login: the default is Username = KEYOP01 (note it's "zero-one" and not "oh-one" at the end); Password = 1313.
# To use the ROI tool built in to the AWS:
Key operator -> System Configuration -> Image Processing Preferences -> Select New View -> Choose view at the top of the page -> Region of Interest
You can then touch on the screen to put a square ROI in place; however, this is quite tricky because
you can't window and level the image; also the size of the ROI is similar to the size of our contrast disc.
What might be easier is to push the image to the PACS and measure it there (Intellerad has an easy-to-use ROI tool).
1) Ensure the CR reader is on and logged in as “Tech” by touching “Tech login” on the screen
2) Open the front door of the reader (below the monitor).
3) Insert the USB drive into one of the slots on the computer case. On the main menu “Eject All USBs” appears as a selection.
4) Select “Image Review” from the main menu.
5) Select “Harvest Mode” from the bottom of the screen.
6) Click on the drive letter beside “Available drive”. Disk type, free space and destination appear.
7) Select “harvest selected images”. An hour glass will appear, and then the Image Review Screen will come up, with check boxes beside each image.
8) Put check marks beside the images you want to put on the USB key.
9) Select “Back”.
10) Select “Compute Needed Space”. An hour glass will come up for a while.
11) Double needed space doesn’t exceed the space on the USB key.
12) Select “Start Harvest”. An hour glass will appear.
13) When “Harvest Accomplished” appears on the screen select “Main Menu”.
14) Select “Eject all USBs”.
15) Remove USB key and close front door of reader.
Image harvesting in the Elite CR should be exactly the same as in the other CR systems.... ..... They removed the PC from the CR reader to make the units smaller and the Elite CR's no longer have a in built UPS. That is the only difference. The software and scanning mechanism is exactly the same, however it is a manual feed system, unlike the other CR units.
Whoops, the only other difference worth mentioning is that the Classic or Elite CR's, which both look exactly the same, but have different throughputs (elite being slightly faster) will only work on version 5 software (which is the new software). The older 850/950's will work with either v4 or v5..... The harvest procedure attached is for v5 (let me know if you need v4)
Siemens
Tips for Siemens machine (digital)
1. tab-del-send to activate and deactivate windows, then you need to hit the “windows key to get the start toolbar, and select it to get the explorer window up.
2. Proc: select "QC RAW" to allowing imaging in raw forma
Hologic
To access modes like getting raw images, you have to log fully out of
the system, then log back in as Physics (Userid= Phys, Pass=Phys). It
takes some time for this to reboot. You have to select for raw images,
and Add procedures under Phys.
Raw images cannot be sent to the RWS.
Images must be burned at the Acq. station CD writer. This is a Unix system.
No images may be uploaded to the Acq. station, so we can't look at our
pattern. They have their own SMPTE, though. Maybe this is something we
need to request from the Lorad people.
Konica
All images sent to PACS are processed, log format.

## Physicist Site Summary
Version 2.54
Full Field Digital Mammography Equipment Evaluation
Unit Number: |  |  | 1 |  |  |  | Room ID | Think Pink
Site: | [6 empty] | Survey Date: | 2005-01-01
Address: | [5 empty] | Report Date: |  | 2005-01-02
Tests Performed By: | [5 empty] | QC Tech. | [5 empty] | CR Reader Manuf.
|  |  | Medical Physicist: |  |  | Signature
Address: |  |  | Sunnybrook HSC, Rm S-632 |  |  | Telephone | [5 empty] | NA
|  |  | 2075 Bayview Avenue, Toronto ON M4N 3M5 | [8 empty] | Agfa
Telephone: |  |  | 416-480-5705 |  |  | Service Eng. | [5 empty] | Fuji
Kodak
|  | Equipment | [9 empty] | Konica
|  |  | X-Ray Unit Manufacturer: |  |  | Model:
|  |  | Installation Date: | 2007-09-28 |  | Serial Num:
|  |  | Detector Serial Num: |  |  | Tube Insert Serial Num:
|  |  | CR Reader Manuf: | 1 |  | Model:
|  |  | Software Version: |  |  | Serial Num:
|  |  | Acquisition Workstation Manuf: |  |  | Model:
|  |  | Software Version: |  |  | Monitor Brand:
|  |  | Review Workstation 1 Manuf: |  |  | Model:
|  |  | Software Version: |  |  | Monitor Brand:
Review Workstation 2 Manuf: | [5 empty] | Model:
|  | Software Version: |  |  |  | Monitor Brand:
|  |  | Laser Printer Manufacturer: |  |  | Model:
Serial Num:
|  |  | Laser Film Manufacturer: |  |  | Type:
|  | Processor Manufacturer: |  |  |  | Model:
| Does the site have: |  |  |  | Y/N
Film Accreditation Phantom
OBSP Dig Mam Uniform Phantom
1 mm discs
Accreditation Status
Accreditation Expiry
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
Full-Field Digital
|  | Site |  |  |  | Report Date | 2005-01-02
[6 empty] | Survey Date | 2005-01-01
|  | X-Ray Unit Manufacturer |  |  |  | Model
|  | Date of Installation |  | 2007-09-28 |  | Room ID | Think Pink
|  | Medical Physicist |  |  |  | Signature
Medical Physicist's QC Tests
1. Mammographic Unit Assembly Evaluation
|  |  | Performs according to 2007 OBSP Digital Mammography Quality Control Manual |  |  |  | H9.2a,d-f, H13b,c
|  |  | Autodecompression can be overidden to maintain compression (& status maintained) | [5 empty] | Pass
|  |  | Manual emergency compression release can be activated in the event of power failure | [5 empty] | Pass
|  | 2. Technique Chart/AEC Evaluation (SDNR) |  |  |  |  | H11.2
SNR and SDNR values for each thickness greater than system-specific minimums
Density Adjustment functions appropriately
|  | 3. Artifact Evaluation |  |  |  |  | H9.3bii
|  |  | Artifacts were not apparent or not significant | [5 empty] | Pass
|  | 4. kVp Accuracy |  |  |  |  | H18.3
Measured kVp within +/-1kVp of nominal kVp
|  | 5. Beam Quality Assessment (HVL Measurement) |  |  |  |  | H9.3biii,H18.4
Half-value layer is within acceptable lower and upper limits at all kV values tested
|  | 6. Breast Entrance Exposure/Average Glandular Dose |  |  |  |  | H8.15
|  |  | Average glandular dose for average breast is below 3.0 mGy (300 mrad) |  |  |  | #VALUE! | mGy | #VALUE!
|  |  | Displayed average glandular is within 20% of physicist calculated value | [7 empty] | This often fails
7. Evaluation of Ghosting and Lag
No evidence of ghosting under clinical viewing conditions.
|  | 8. Collimation Assessment |  |  |  |  | H9.2b,H18.1a,H18.2
Field Indicator > 160 Lux above Background
Light-field and Image Receptor edges with 10 mm
X-Ray field does not extend beyond image Image Receptor by more than 13 mm
|  |  | Compression Paddle does not extend beyond Image Receptor by more than 5 mm | [5 empty] | Pass
Amount of missed tissue less than 7 mm
9. Evaluation of System Resolution
MTF higher than OBSP reference values for this system type
MTF at all frequencies within 10% of acceptance value
|  | 10. Noise and Linearity |  |  |  |  | H13.5
Detector Signal Uniformity - ADU values within 20% of central value
|  |  | Detector Noise Uniformity - SNR values within 30% of central value | [5 empty] | #VALUE!
|  |  | Detector Linearity - all values within 10% of mean | [7 empty] | Linearity may fail if lowest mAs too low?
Noise Linearity - fit to variance vs exposure has R2 greater than 0.95
11. Spatial Linearity and Geometric Distortion Evaluation
Image shows uniform resolution and straight lines with minimal distortion.
Effective pixel width and length are within 5% of each other
Image within 5% of manufacturer's stated nominal image size
Version 2.54
MEDICAL PHYSICIST'S MAMMOGRAPHY QC TEST SUMMARY
Hard and Soft-Copy Display for Full-Field Digital
|  | Site |  |  |  | Report Date | 2005-01-02
[6 empty] | Survey Date | 2005-01-01
|  | X-Ray Unit Manufacturer |  |  |  | Model
|  | Date of Installation |  | 2007-09-28 |  | Room ID | Think Pink
|  | Review Workstation Manuf. |  |  |  | Model
|  | Review Workstation Manuf. |  |  |  | Model
|  | Printer Manufacturer |  |  |  | Model
|  | Medical Physicist |  |  |  | Signature
Medical Physicist's QC Tests
12. Monitor Display Quality Evaluation
Acquistion Monitor TG18-QC Pattern Evaluation
|  |  | Review Workstation TG18-QC Pattern Evaluation | [5 empty] | Pass
Review Workstation distance measurement tool accurate to 5%
Monitor Uniformity
13. Review Monitor Luminance Response Measurement and Viewing Conditions Assessment
|  |  | Maximum intensity is at least 171 cd/m2 | [8 empty] | Max intensity limit | 171
Maximum intensities within 10%
Dynamic range at least 250:1
Response within 15% of DICOM gray-scale display function
JNDs per luminance interval meet 2007 OBSP Digital Mammography Quality Control Manual
Room illumination level below 10 lux
14. Viewbox Luminance and Room Illuminance
Viewbox Luminance ≥ 3000 candela per square meter
Room illumination level below 10 lux
15. Laser Printer Evaluation
TG18-QC Pattern is acceptable
|  |  | Dynamic range at least 100:1 | [8 empty] | 100 | Dynamic Range Limit
|  |  | Printer response is within 20% of DICOM gray-scale display function | [8 empty] | 20% | GSDF Limit
JNDs per luminance interval meet 2006 ACR Digital Mammography Quality Control Manual
No noticeable/objectionable artifacts
|  | 16. Technologist's QC Program Evaluation |  |  |  |  | H8.10
|  |  | QC program as specified in 2007 OBSP Digital Mammography Quality Control Manual | [5 empty] | Pass
17. CR Interplate Sensitivity
18x24 mAs within +/-10% of average
24x30 mAs within +/-10% of average
18x24 SNR within +/-15% of average
24x30 SNR within +/-15% of average
18. Kodak CR Cassette Exposure Response Test
EI within 100 units of target value
18. Fuji CR Sensitivity Check
S Number within 15% of target value
18. Konica CR Sensitivity Check
Relative Exposure Number within 20% of target value
Physicist's Comments
Version 2.54
Note on E20: Aili K Bloomquist: Dicom Tag 0x0018,0x700A Not applicable for CR.

## Unit Evaluation 1
2005-01-01
Chart 1 - Full Field Digital Mammography Equipment Evaluation | [6 empty] | Pass
at acquisition station
Full Field Digital Mammographic Unit Assembly Evaluation | [8 empty] | at review station
automatic
1. Free standing unit is mechanically stable. | [6 empty] | 1
[10 empty] | Pass | 1
2. All moving parts move smoothly, without obstructions to motion. | [6 empty] | 1 |  |  | Fail | 2
[10 empty] | Marginal | 3
3. All locks and detents work properly. | [6 empty] | 1 |  |  | N/A | 4
[10 empty] | Not Done | 5
4. Image receptor is free from vibrations during exposure. | [6 empty] | 1
5. Compressed breast thickness scale is accurate to +/- 5 mm, reproducible to +/- 2 mm. | [9 empty] | Compressed breast thickness accuracy: |  |  |  | 5
Displayed Thickness (mm)
Actual Thickness (mm)
Enter "P" or "F" in each yellow box.
6. Automatic decompression override exists (if applicable) and override status displayed. | [6 empty] | 1
7. Manual compression release possible with no power to machine. | [6 empty] | 1
8. Patient or operator is not exposed to sharp or rough edges or other hazards. | [6 empty] | 1
9. Operator protected during exposure by adequate radiation shielding. | [6 empty] | 1
4.0 cm phantom compressed to 8-10 daN
10. Operator technique charts are posted. | [6 empty] | 1
11. Indicator lights working properly. | [6 empty] | 1
12. Maximum Compression pressure applied under power less than 20 daN? | [9 empty] | Max compression force: |  |  | 20 | daN
[13 empty] | 45 | lbs
|  | lbs | max 45
|  | daN | max 20
13. Back-up timer functioning correctly | [6 empty] | 1
14. Annotation on images correct. | [6 empty] | 1
15. DICOM header correctly populated.
Other:
Version 2.54
2005-01-01
Chart 1-B Dicom Image Header Checklist
Required
| Tag Number |  | Tag Description | Present | Correct | DICOM | OBSP |  |  | Pass/Fail
| 0008,0080 |  | Institution Name | 3 | 3 | N | Y |  |  |  |  | Y | 1
| 0008,0022 |  | Acquisition Date | 3 | 3 | N | Y |  |  |  |  | N | 2
| 0008,0032 |  | Acquisition Time | 3 | 3 | N | N | [5 empty] | 3
| 0008,0081 |  | Institution Address | 3 | 3 | N | Y*
| 0008,1010 |  | Station Name | 3 | 3 | N | Y^
| 0008,1070 |  | Operator | 3 | 3 | N | Y
| 0010,0010 |  | Patient Name | 3 | 3 | N | Y
| 0010,0020 |  | Patient ID | 3 | 3 | N | Y
| 0018,0060 |  | kV | 3 | 3 | N | N
| 0018,1150 |  | Exposure Time | 3 | 3 | N | N
| 0018,1151 |  | X-ray Tube Current | 3 | 3 | N | N
| 0018,1152 |  | Exposure | 3 | 3 | N | N
| 0018,1191 |  | Anode Target Material | 3 | 3 | N | N
| 0018,5101 |  | View Position | 3 | 3 | N | Y
| 0018,7050 |  | Filter Material | 3 | 3 | N | N
| 0020,0020 |  | Patient Orientation | 3 | 3 | Y | N |  | CR this might be 18,5100 and 18,5101
| 0020,0062 |  | Image Laterality | 3 | 3 | Y | Y
Overall Pass /Fail
*The address shall include at least the city, province and postal code
of the institution.
^Identification of the unit is required if the site has multiple units.
Version 2.54
Note on B63: Aili K Bloomquist: Tag 54,220 View Code sequence IS mandatory in dicom, and the "View" can be deciphered from it (with effort), so perhaps we should make this optional
Note on D65: Aili K Bloomquist: CR this might be 18,5100 and 18,5101
Note on D66: Aili K Bloomquist: CR this might be 20,60
Pass DICOM

## A-Image Acquisition
Table A - Raw Image Acquistion | [23 empty] | suggested patient name: "Physics, Testing"
automatic
| Test | Dose / Baseline | Artifact Evaluation (Tabletop) |  |  |  | MTF |  |  |  | Artifact Evaluation (Mag) |  | Thickness Tracking |  |  |  | AEC Density Steps (if applicable) | [7 empty] | at acquisition station
| Phantom | DMUP with disc | Site's DMUP Phantom, with Cont. Disc |  |  |  | MTF Tool or OBSP Digital Physics Phantom |  |  |  | 2 cm equiv |  | Uniform Phantoms 2, 4, 6 and 8 cm equiv., with Cont. Disc and post for Al. |  |  |  | 4 cm Equivalent Uniform Phantom and post if Aluminum. | [6 empty] | For acceptance testing, don't forget to evaluate 2 and 6 cm of PMMA or 1 mm and 3 mm Aluminum for artifacts as well.
| Material | PMMA | PMMA |  |  |  | PMMA + overlay | PMMA + overlay | PMMA + overlay | PMMA + overlay | Al | Al | Al | Al | Al | Al | Aluminum | [6 empty] | For Digital Accred. Phantom image using all anodes and focal spot sizes
| Thickness (mm) | 40 | 40 |  |  |  | 40 | 40 | 40 | 40 | 1.27 | 1.27 | 1.27 | 2.54 | 3.8 | 5.1 | 2 | [6 empty] | When using contrast disc, place disc approx 5 cm from chest wall edge, centered left-right in field
| Technique | AEC | Anode 1 Filter 1 | Anode 1 Filter 2 | Anode 1 Filter 3 | Anode 2 Filter 2 | Anode 1 | Anode 2 | Anode 1 Mag | Anode 2 Mag | Mag 1 Filter 1 | Mag 1 Filter 2 | AEC | AEC | AEC | AEC | -5 | -3 | -1 | 0 | 1 | 3 | 5 | Currently using "4 cm" exposure from thickness tracking for dose and standard exposure.
Lead # for CR
CR Processing Menu
| AEC mode | AEC Auto | man | man | man | man | man | man |  |  | AEC Auto | man |  |  |  |  | man | man | man | man | man | man | man
| AEC Sensor Position | [22 empty] | For Ghosting wait 1 minute between acquisitions.
| Target | [22 empty] | For collimation use all anodes and FOV sizes. Fill in Chart 10 at acquisition time if phosphor screen glows |  | Do we bother with both Anodes for non-mag?
Filter
| Focal Spot Size | Large | Large | Large | Large | Large | Large | Large | Small | Small | Small | Small | Large | Large | Large | Large | Large | Large | Large | Large | Large | Large | Large
kV Setting
mAs
| S, EI or Rel Exp (CR) | [22 empty] | Don't forget to measure light field intensity for both collimator sizes.
L Number for CR
| Displayed Thickness | [22 empty] | Don't forget to measure kV min and max, and HVL/tube output.
Displayed Comp. Force
Acquisition Time
Window width
Window level
ROI1
ROI2
St Dev
Displayed Dose
| Magnification Factor | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  |  |  |  | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
Comments
For CR Systems check interplate sensitivity.
Table A - Raw Image Acquistion - Continued | [26 empty] | For Fuji CR systems very sensitivity calibration
| Test | Noise and Linearity |  |  |  | Ghosting |  | Collimation and Distortion |  |  |  | CR Interplate Senstivity - See Chart 18 | Kodak CR Cassette Response - See Chart 19 | Fuji CR Sensitivity - See Chart 19 | BUT |  |  | For CR checkdistortion on both cassette sizes.
| Phantom | DMUP with contrast disc |  |  |  | DMUP |  | Radio Opaque Rulers, Circuit board, Phosphorescent Screen | [6 empty] | Beamstop
| Material | PMMA |  |  |  | PMMA |  | NA | NA | NA | NA |  |  |  | Lead or SS Sheet
| Thickness (mm) | 40 mm |  |  |  | 40 | 40 | NA | NA | NA | NA |  |  |  | 2
| Technique | low mAs | mid mAs | Standard mAs | high mAs | Man, phant. to rt | Man, phant. centred | Small FOV | Small Left | Small Rt. | Large FOV |  |  |  | AEC
Lead # for CR
CR Processing Menu
| AEC mode | man | man | man | man | man | man | man | man | man | man | AEC | man | man | man
| AEC Sensor Position |  |  |  |  | Collimator -> |  | 18x24 | 18x24 L | 18x24 R | 24x30 | [6 empty] | Process with "Physics, Sensitivity"
| Target | [10 empty] | Mo | Mo | Mo
| Filter | [10 empty] | Mo | Mo | Mo
| Focal Spot Size | Large | Large | Large | Large | Large | Large | Large | Large | Large | Large | Large | Large | Large | Large
| kV Setting | [10 empty] | 28 | 28 | 25
mAs
S or EI Number for CR
L Number for CR
Light Field intensity LUX
1
Note on B57: Aili Bloomquist: EI is tag 18,1405 S is tag 18,6000
Note on D7: Aili K. Bloomquist: Do all allowed target-filter combinations
Note on D56: Gord Mawdsley: mAs closest to standard
Note on H8: Gord Mawdsley: This image is not ued fro dose because the overlay might affect the AEC operation
Note on H12: Gord Mawdsley: GE- Probably Moly
Note on H13: Gord Mawdsley: GE- Probably Moly
Note on H15: Gord Mawdsley: GE- Probably Moly
Note on H16: Gord Mawdsley: GE Probably Rh
Note on H18: Gord Mawdsley: GE Probably Rh
Note on H19: Gord Mawdsley: GE should give about 56 mAs
Note on I15: Gord Mawdsley: Use other target
Note on I16: Gord Mawdsley: Use other filter
Note on K15: Aili K. Bloomquist: Use other Target
Note on P43: Aili K Bloomquist: Test of Back-Up Timer Function. BUT must kick in or exposure must abort, with reset required.
Artifact Evaluation using Site's phantom
It is not necessary to image ALL permutations with the site's phantom for artifact eval. mAs setting should be similar to that used for patients. Mo, Rh - mAs will be similar. W may be higher, but it is not critical.
Use a 2 cm equiv. phantom to acquire images on the Mag stand at clinically used height and anode/filter combos.
Use 1/8 x,1/2 x,1x and 4x the mAs used to image the 4 cm uniform phantom for the dose test. Use closest value and type over the formula value.
AEC Density steps are evaluated on CR systems. All that is needed is the mAs chosen for each setting. Don't process the plates (use the same plate over and then erase it afterwards).

## kVp 4 and HVL 5
2005-01-01
Ensure one of the kV, Target and Filter combinations matches that used for the standard breast
Chart 5 - Tube Output, Beam Quality (HVL) Measurement and kVp Accuracy | [9 empty] | exposure and MGD image (Test 7)
|  |  | Dosimetry system used: |  |  | kVp meter used: |  |  |  |  | Typical HVLs |  |  |  | Error Limit for set kV>= |  | 35 | 0.04 | *kV
|  |  | Imaging receptor size: |  |  |  | Dosimeter Corr. Factor | 1 |  |  | Mo/Mo | 26 | 0.32 |  | Error Limit for set kV < 35 |  |  | 1
|  |  | SID (mm) | 650 | [6 empty] | Mo/Rh | 27 | 0.4
[11 empty] | Rh/Rh | 27 | 0.45
|  |  |  | Nominal | Measured | Pass/Fail |  |  |  |  | W/Rh | 27 | 0.54
Lowest Clinical kVp
Common Clinical kVp
Highest Clinical kVp
| N/A |  | HVL for / (behind compressor at breast surface) | [6 empty] | Protect the detector using a lead sheet.
|  |  | mm Aluminum | mR | Ratio | [5 empty] | target mR |  | Note calculation of HVL occurs if 1st data row is completed
|  | kV | 0 |  | 0 | Target |  | Filter |  |  |  |  | and two subsequent rows bracket 0.5*mR at 0 mm.
| 50 | mAs (mA) | 0.325 |  | 0 |  | Minimum HVL:
|  | C | 0.35 |  | 0 |  | Maximum HVL: | [5 empty] | Note also that calculation REQUIRES mR values to bracket 0.5*mR for 0 mm Al.
|  | mR/mAs | 0.375 |  | 0 |  | / HVL = | #VALUE!
|  |  | 0 |  | 0
[10 empty] | Discard images |  |  |  | fill in while at acquisition station/acquiring images
| N/A |  | HVL for / (behind compressor at breast surface)
|  |  | mm Aluminum | mR | Ratio | [8 empty] | Filled in by the software
|  | kV | 0 |  | 0 | Target |  | Filter
| 50 | mAs (mA) | 0.375 |  | 0 |  | Minimum HVL:
|  | C | 0.425 |  | 0 |  | Maximum HVL:
|  | mR/mAs | 0.45 |  | 0 |  | / HVL = | #VALUE!
|  |  | 0 |  | 0
| N/A |  | HVL for / (behind compressor at breast surface)
|  |  | mm Aluminum | mR | Ratio
|  | kV | 0 |  | 0 | Target |  | Filter
| 50 | mAs (mA) | 0.425 |  | 0 |  | Minimum HVL:
|  | C | 0.45 |  | 0 |  | Maximum HVL:
|  | mR/mAs | 0.475 |  | 0 |  | / HVL = | #VALUE!
|  |  | 0 |  | 0
| N/A |  | HVL for / (behind compressor at breast surface)
|  |  | mm Aluminum | mR | Ratio
|  | kV | 0 |  | 0 | Target |  | Filter
| 50 | mAs (mA) | 0.425 |  | 0 |  | Minimum HVL:
|  | C | 0.45 |  | 0 |  | Maximum HVL:
|  | mR/mAs | 0.475 |  | 0 |  | / HVL = | #VALUE!
|  |  | 0 |  | 0
close exam and delete images.
|  |  | Tube Reproducibility - repeated exposures within 2% | [7 empty] | 2% | Reproducibility limit which is considered equivalent to COV of | [5 empty] | 0.05 | for 5 measurements
[12 empty] | and to having all measurements within |  |  |  | 15% | of mean
Comments
Version 2.54
|  |  |  | Lookup for Definitions of "C" | [6 empty] | Lookup for Definitions of "C"
[11 empty] | Target/Filter | C
|  |  |  | 0.12 for Mo/Mo |  |  | 0.28 for W/Mo |  |  |  | Mo/Mo | 0.12
|  |  |  | 0.19 for Mo/Rh |  |  | 0.30 for W/Rh |  |  |  | Mo/Rh | 0.19
|  |  |  | 0.22 for Rh/Rh |  |  | 0.32 for W/Al |  |  |  | Rh/Rh | 0.22
[11 empty] | W/Mo | 0.28
|  |  |  | Note, these standards are tighter than MQSA | [6 empty] | W/Rh | 0.3
[11 empty] | W/Al | 0.32
Note on C22: Gord Mawdsley: Ues a different filter from the previous test
Protect the detector with a lead sheet if possible
Remember to do the unfiltered (0 mm Aluminum) entrance exposure measurements TWICE to check tube reproducibility! Delete this box before submitting report.

## Collimation 8
2005-01-01
Chart 8- X-Ray Field Evaluation | [11 empty] | Enter deviation in mm or "Y" if less than 10 mm and "N" if greater than 10 mm
Suggested Patient Name: "Physics, Collim"
|  | Light Field Intensity | [5 empty] | Lux |  |  |  | Suggested ID number 9906UYYMMDD |  |  | For Fuji CR process using "Auto Mode" with a fixed latitude of 2.0
|  | Collimator |  |  | 18x24 | 24x30 |  | 18x24 | 24x30 | Bckgnd
|  | Pass/Fail |  |  | 5 | 5 | [5 empty] | Pass | 1 |  | fill in while at acquisition station/acquiring images
| Action Limit: | Field Intensity must be 160 lux above background | [9 empty] | Fail | 2
[12 empty] | Marginal | 3 |  | fill in while at review station/using software to analyze images
| Mark edges of light field and paddle with radiographic rulers. | [10 empty] | NA | 4 |  | Filled in by the software
5
|  | Light Field - Detetector Congruence: | [12 empty] | Lux Intensity Limit: | 160
|  | Collimator |  |  | 18x24 | 18x24 L | 18x24 R | 24x30 |  |  |  | 26 Kvp 80 mAs Mo/Mo works well for GE |  |  | Light-Field Limit: | 10
|  | Anode | [9 empty] | Enter "Y", "N" or deviation in mm
|  | Left Edge: |  | Deviation < 10 mm?
|  | Right Edge: |  | Deviation < 10 mm?
|  | Anterior Edge: |  | Deviation < 10 mm?
|  | Chest Edge: |  | Deviation < 10 mm? | [12 empty] | fill in while at acquisition station/acquiring images
fill in while at review station/using software to analyze images
| Action Limit: | If the light field and detector edges differ by more than 10 mm, seek service adjustment. | [14 empty] | Filled in by the software
|  | X-ray field exceeds limits: | [9 empty] | Enter "Y", "N" or deviation in mm |  |  | Pass
|  | Collimator |  |  | 18x24 | 18x24 L | 18x24 R | 24x30
|  | Anode | [12 empty] | X-Ray field limit: | 13
|  | Left Edge: |  | Deviation <= 13 mm? | [10 empty] | CW X-Ray field limit: | 7
|  | Right Edge: |  | Deviation <= 13 mm?
|  | Anterior Edge |  | Deviation <= 13 mm?
|  | Chest Edge |  | Deviation <= 7 mm?
| Action Limit: | If the radiation field extends more than 13 mm from left, right or anterior edge,
or more than 7 mm beyond chest wall edge, seek service adjustment.
[12 empty] | Missed Tissue limit: |  | 7
Missing Tissue (FFDM Phantom)
|  |  | left scale (mm) |  |  |  | right scale (mm)
| Action Limit: | If the missed tissue exceeds 7 mm for either scale, seek service adjustment.
| Pass | Coverage of Image detector | [9 empty] | deviation between x-ray field and image receptor made by
|  | Collimator |  |  | 18x24 | 18x24 L | 18x24 R | 24x30 |  |  |  | comparing film and digital images.
Anode
|  | Entire Active Area Covered? |  |  | 5 | 5 | 5 | 5
| Action Limit: | If any edge of the x-ray field is visible in the image, seek service adjustment.
| Pass | Alignment of chest wall edges of compression paddle and image receptor:
|  | Collimator |  |  | 18x24 | 18x24 L | 18x24 R | 24x30
|  | Anode | [9 empty] | Comp. Paddle - Receptor limit: | 5
|  | Paddle Edge Not Visible? (Y/N) |  |  | 5 | 5 | 5 | 5
Difference between compression paddle edge and image receptor at chest wall < 5 mm?
| Action Limit: | If chest wall edge of compression paddle is within the image receptor or projects beyond | [9 empty] | close exam
|  | the chest wall edge of the image receptor by more than 5 mm, seek service correction. | [8 empty] | Version 2.54
Note on E18: Aili Bloomquist: Should this be Deviation < 13 mm or field within bucky?
Note on E28: Aili Bloomquist: Should this be Deviation < 13 mm or field within bucky?

## Acq. Display Quality 12 A
2005-01-01
Chart 12-A - Acquisition Monitor TG18-QC Pattern and Uniformity Evaluation
| Use TG18-QC Image |  |  |  |  | Note: If two sizes of TG18 pattern are available, then use the larger size for this test.
Workstation:
General Image Quality
| Smearing Absent | 5 |  |  |  |  | fill in while at review station/using software to analyze images
| Artifacts Absent | 5 |  |  |  |  | Filled in by the software
| Ramps continuous | 5
| Geometric Distortion | [5 empty] | Pass | 1
| Lines straight | 5 |  |  |  |  | Fail | 2
| Pattern centered | 5 |  |  |  |  | Marginal | 3
| Boxes square | 5 |  |  |  |  | NA | 4
| Luminance | [6 empty] | 5
| Patches distinct | 5
| Patch corners visible | 5
| 0-5% | 5 |  |  |  | Letter Threshold: |  | 10
| 95-100% | 5
| Lettering (# visible) |  | Minimum - Quality Con = 10 |  |  | 14 | QUALITY CONTROL
| Dark |  |  |  |  | 13 | QUALITY CONTRO
| Mid-Gray |  |  |  |  | 12 | QUALITY CONTR
| Light |  |  |  |  | 11 | QUALITY CONT
[6 empty] | 10 | QUALITY CON
| Overall Pass/Fail |  |  |  |  | 9 | QUALITY CO
[6 empty] | 8 | QUALITY C
| Comments: |  |  |  |  | 7 | QUALITY
| For artifact evaluation enter "P" or "F" for pass or fail. |  |  |  | Version 2.54

## Rev. Display Quality 12 B
2005-01-01
Chart 12 B - Review Monitor TG18-QC Pattern and Uniformity Evaluation
| Use TG18-QC Image | [11 empty] | P | 1 | P | 1
| Workstation: |  |  |  |  | All Answers "Y" or "N" | [6 empty] | F | 2 | F | 2
[13 empty] | M | 3 | M | 3
| Monitor Similarity | P | [10 empty] | NA | 4 |  | 4
| Right and Left Monitors Same? | 1 | [11 empty] | 5
[10 empty] | Hor. | Ver.
|  | Monitor |  |  | Measured 50 mm Ruler Length:
|  | L | R |  | Accurate to +/- 5%
| General Image Quality | P | P
| Smearing Absent | 1 | 1 |  |  |  |  | Sm. Img. |  | Lg. Img. |  |  |  | fill in while at review station/using software to analyze images
| Artifacts Absent | 1 | 1 |  |  |  |  | Monitor |  | Monitor |  |  |  | Filled in by the software
| Ramps continuous | 1 | 1 |  |  |  |  | L | R | L | R
| Geometric Distortion | P | P |  | Monitor Resolution |  |  | P | P | P | P |  | Examine resolution lines with images set to "Full Resolution" (zoom of 1)
| Lines straight | 1 | 1 |  | Nyquist Lines Visible:
| Pattern centered | 1 | 1 |  | Top Left: |  | Horiz. | 1 | 1 | 1 | 1 |  | 14 | QUALITY CONTROL
| Boxes square | 1 | 1 |  |  |  | Vert. | 1 | 1 | 1 | 1 |  | 13 | QUALITY CONTRO
| Luminance | P | P |  | Top Right: |  | Horiz. | 1 | 1 | 1 | 1 |  | 12 | QUALITY CONTR
| Patches distinct | 1 | 1 |  |  |  | Vert. | 1 | 1 | 1 | 1 |  | 11 | QUALITY CONT
| Patch corners visible | 1 | 1 |  | Middle: |  | Horiz. | 1 | 1 | 1 | 1 |  | 10 | QUALITY CON
| 0-5% | 1 | 1 |  |  |  | Vert. | 1 | 1 | 1 | 1 |  | 9 | QUALITY CO
| 95-100% | 1 | 1 |  | Bottom L: |  | Horiz. | 1 | 1 | 1 | 1 |  | 8 | QUALITY C
| Lettering (# visible) | [5 empty] | Vert. | 1 | 1 | 1 | 1 |  | 7 | QUALITY
| Dark |  |  |  | Bottom R: |  | Horiz. | 1 | 1 | 1 | 1
| Mid-Gray | [5 empty] | Vert. | 1 | 1 | 1 | 1
Light
| Overall Pass/Fail | Pass
Uniformity Evaluation | [12 empty] | Ruler accuracy limit |  | 5%
| Use UNL20 and UNL 80 Images | [11 empty] | Letter limit |  | 11
|  | Luminance | [5 empty] | Measurements should be |  |  |  |  | Uniformity limit |  | 30%
|  | Left |  | Right |  |  |  | within 30% of average value
| Location | 20 | 80 | 20 |  | 80
Top Left - nit
Top Right
Middle
Bottom Left
Bottom Right
| Artifacts Absent | 4 | 4 | 4 | 4 | 4
| Uniformity Pass/Fail | [10 empty] | Version 2.54
Note on H19: Aili Bloomquist: View Images at 1:1 res. For this portion

## Monitor Lum. Response 13
2005-01-01
fill in while at review station/using software to analyze images
Chart 13 - Monitor Luminance Response | [9 empty] | Filled in by the software
UseTG18-LN Images
| Workstation: | [7 empty] | Ambient light limit |  | 10
[9 empty] | Monitor matching limit |  | 10%
|  | Ambient Luminance | [6 empty] | Luminance Ratios limit |  | 250
|  | Left | Right | [5 empty] | GSDF limit |  | 15%
Nits
[12 empty] | Ideal JNDs Left | Ideal JNDs Right | Background pix val: |  | 153 | [9 empty] | Background
| Target | Luminance (in Nits) | [5 empty] | L' Left (L + Ambient) | L' Right (L + Ambient) | Measured |  |  |  | Expected Luminance Left | Expected Luminance Right | Normalized Contrast Left | Normalized Contrast Right | Norm. Exp. Contrast Left | Norm. Exp. Contrast Right | Pass/ Fail Left | Pass/ Fail Right | lower limit Left | lower limit Right | upper limit Left | upper limit Right | 8 Bit |  |  | 12 Bit
|  | Left | Right | [23 empty] | Deviation Left | Deviation Right
Min | 1 | [8 empty] | JNDs Left | JNDs Right | [14 empty] | Pix vals
| 2 | [24 empty] | 0 |  |  | 0
| 3 | [24 empty] | 15 |  |  | 241
| 4 |  |  |  | Only include ambient luminance | [20 empty] | 30 |  |  | 482
| 5 |  |  |  | if Luminance measurements | [20 empty] | 45 |  |  | 723
| 6 |  |  |  | are taken with a meter which is | [20 empty] | 60 |  |  | 964
| 7 |  |  |  | used by putting it very gently against | [20 empty] | 75 |  |  | 1204
| 8 |  |  |  | the monitor surface. Telescopic | [20 empty] | 90 |  |  | 1445
| 9 |  |  |  | photometer measurements | [20 empty] | 105 |  |  | 1686
| 10 |  |  |  | already include the ambient | [20 empty] | 120 |  |  | 1927
| 11 |  |  |  | luminance information. | [20 empty] | 135 |  |  | 2168
| 12 |  |  |  | To obtain ambient illumination | [20 empty] | 150 |  |  | 2409
| 13 |  |  |  | hold the photometer 50 cm | [20 empty] | 165 |  |  | 2650
| 14 |  |  |  | away from the monitor surface, | [20 empty] | 180 |  |  | 2891
| 15 |  |  |  | pointing towards the monitor, | [20 empty] | 195 |  |  | 3131
| 16 |  |  |  | with the monitor turned off. | [20 empty] | 210 |  |  | 3372
| 17 | [24 empty] | 225 |  |  | 3613
Max | 18 | [24 empty] | 240 |  |  | 3854
[26 empty] | 255 |  |  | 4095
| Ambient Illuminance |  | Lux
| Ambient Light Low |  | Lamb<10 lux | [18 empty] | 15
| Monitors L'max matched |  | L'max values within 10%
| Luminance Ratios |  | L'max/L'min > 250
| Response |  | within 15% of GSDF
Version 2.54
2005-01-01
Chart 13-B - Monitor Luminance Response
To Calculate JND index as a function of Luminance
j(L) = A + B*Log10(L) + C*(Log10(L))2 + D*(Log10(L))3 + E*(Log10(L))4 + F*(Log10(L))5 + G*(Log10(L))6 + H*(Log10(L))7 + I*(Log10(L))8
[8 empty] | A |  | 71.4981
[8 empty] | B |  | 94.5931
[8 empty] | C |  | 41.9121
[8 empty] | D |  | 9.8247
[8 empty] | E |  | 0.28175
[8 empty] | F |  | -1.1878
[8 empty] | G |  | -0.18014
[8 empty] | H |  | 0.14711
[8 empty] | I |  | -0.017047
To Calculate Luminance as a function of JND index
log10L(j) = (a + c*Ln(j) + e*(Ln(j))2 + g*(Ln(j))3 + m*(Ln(j))4)/(1 + b*Ln(j) + d*(Ln(j))2 + f*(Ln(j))3 + h*(Ln(j))4 + k*(Ln(j))5)
[8 empty] | a |  | -1.3012
[8 empty] | b |  | -0.02584
[8 empty] | c |  | 0.080243
[8 empty] | d |  | -0.1032
[8 empty] | e |  | 0.13647
[8 empty] | f |  | 0.028746
[8 empty] | g |  | -0.025468
[8 empty] | h |  | -0.0031979
[8 empty] | k |  | 0.00012993
[8 empty] | m |  | 0.0013635
| Metric |  | Left | Right |  | Pass/Fail | Criteria |  |  | Criteria
| Jp (ideal JNDs/interval) | [5 empty] | < 3 |  |  | 3
| Max. Dev. from Jp | [5 empty] | < 2 |  |  | 2
| RMS Dev. from Jp | [5 empty] | < 1 |  |  | 1
Version 2.54
Note on B84: Aili K Bloomquist: These equations assume 8 bit display board.

## Monitor View. Conditions 13-C
2005-01-01
Chart 13-C - Acceptable Viewing Conditions for Workstation Monitors
fill in while at review station/using software to analyze images
To maintain sufficiently low ambient light levels for viewing digital mammograms,
the following room/lighting configuration should be used:
Suggested Acceptable conditions
| Other viewboxes in room turned off, or fully masked. |  | Viewboxes masked
| Room fluorescent lights turned off. |  | Overhead lights off
| Hallway door closed. |  | Blinds drawn
| Blackout blinds fully drawn. |  | Doors closed
Version 2.54

## Viewbox View. Conditions 14
2005-01-01
Chart 14 - Viewbox Luminance and Room Illuminance
fill in while at review station/using software to analyze images
Filled in by the software
| Ambient Luminance of Viewbox |  | nits |  |  | Limits | Pass | 1
| Viewbox Luminance |  | nits | (min 3000 nits) |  | 3000 | Fail | 2
| Viewbox Uniformity | 1 | (Pass/Fail) |  |  |  |  | 3
| Viewbox Masking | 1 | (Pass/Fail)
| Viewbox Cleanliness | 1 | (Pass/Fail)
| Viewbox Illuminance |  | lux | (max 10 lux) |  | 10
| Overall Pass/Fail | Pass
Version 2.54

## Technologist QC Prog 16
|  |  | 2005-01-01 | 2005-01-01
Chart 16 Technologist's QC Program Evaluation
Pass
| Test | Minimum Frequency | Performed Correctly |  | Enter "Y" or "N" for each item
| Daily Checklist | Daily | 1
fill in while at review station/using software to analyze images
| Phantom Image Quality | Weekly | 1
| Display Monitor QC | Weekly | 1 |  | Pass | 1
[5 empty] | Fail | 2
| CR Full-Field Artifact Evaluation | Monthly | 4 |  | Marginal | 3
| Resolution / MTF Evaluation | Monthly | 4 |  | NA | 4
| Checklist of Exam Room | Monthly | 1 |  |  | 5
| Repeat Analysis | Quarterly | 1
| Compression Force | Semi-Annually | 1
Laser Printer Tests - if images are read from film.
| Test | Minimum Frequency | Performed Correctly
| Laser Printer Sensitometry Check | Daily/Monthly | 4
| Darkroom / Printer Area Cleanliness | Daily | 4
| Viewbox Cleanliness | Weekly | 4
| Laser Printer Artifacts | Monthly | 4
| Printed image Quality Test | Quarterly | 4
| Analysis of Fixer Retention | Quarterly | 4
| Darkroom Fog | Semi-annually | 4
Comments
Version 2.54
Note on B13: Aili K Bloomquist: This test is for systems with moving parts (ie CR) only!

## CR Interplate Sensitivity 17
[7 empty] | 2005-01-01 |  |  | at acquisition station
at review station
17 - CR Interplate Sensitivity Comparison | [9 empty] | automatic
Use equivalent of 4 cm lucite or breast equivalent material or 2 mm Al, sufficient to cover most of the detector.
|  | Patient Name: | [8 empty] | Use a consistent interval between plate exposure and reading (Kodak suggests 5 minutes)
Patient ID:
Phantom Material
Phantom Thickness (mm)
Kodak Suggests 28 kV Mo/Mo with
|  | AEC Mode | AEC | [7 empty] | 2 mm Al as the phantom
Sensor Position
|  | Target | Mo
|  | Filter | Mo
|  | kVp | 28
|  |  | 18x24 | 24x30
mAs within +/- 10%
SNR with +/- 15%
EIs within +/- 50
|  | Plate Size: | 18x24
| Film ID | Imaging Plate ID | mAs | EI or S | Signal | Std. Dev. | SNR | Artefacts (Y/N) |  |  | mAs P/F | mAs P/F | SNR P/F | SNR P/F | EI P/F | EI P/F
|  | Plate Size: | 24x30
| Film ID | Imaging Plate ID | mAs | EI or S | Signal | Std. Dev. | SNR | Artefacts (Y/N)
Version 2.54
|  |  |  | Lower Limit | Upper Limit
|  | 18x24 Average mAs |  |  |  | +/-mAs Limit |  | 10%
|  | 18x24 Average SNR |  |  |  | +/- SNR Limit |  | 15%
|  | 24x30 Average mAs |  |  |  |  | EI Limit | 50
24x30 Average SNR
Note on C22: Aili K Bloomquist: Kodak Only
Note on E25: Aili K. Bloomquist: Kodak Only
Note on E41: Aili K. Bloomquist: Kodak Only

## Kodak CR Exposure Resp. 18
[5 empty] | 2005-01-01 |  | at acquisition station
at review station
18 - Cassette Exposure Response Test - Kodak CR Only | [6 empty] | automatic
Use equivalent 4cm 50/50, sufficient to cover most of the detector.
| Patient Name: | [6 empty] | Use a consistent interval between plate exposure and reading (Kodak suggests 5 minutes)
Patient ID:
[8 empty] | EI Target limit |  | 100
| Phantom Material |  | 50/50
| Phantom Thickness (mm) |  | 40
Kodak says 28 kV Mo/Mo with
| AEC Mode | man | [5 empty] | 2 mm Al as the phantom (raised up near tube port), and
| Target | Mo | [5 empty] | place the cassettes on the tabletop, not in bucky.
| Filter | Mo | [5 empty] | Choose mAs to give ~100 mR exposure behind Al.
| kVp | 28
| mAs | [6 empty] | For Kodak, Expected EI calculated as 1000Xlogbase10(mR)+Beam Quality Offset
| Exposure to Cassette(mR) | [6 empty] | Material | kV | Target | Filter | Beam Quality Offset
| Beam Quality Offset | 1080 | [5 empty] | 2 mm Al | 28 | Mo | Mo | 1000
| Target EI | [6 empty] | 2 cm fat | 25 | Mo | Mo | 1025
[8 empty] | 4 cm 50/50 | 27 | Mo | Mo | 1080
| Casette Size | 18 x 24 cm |  | 24 x 30 cm |  |  |  |  | 27 | Mo | Rh | 1100
| Imaging Plate ID | [7 empty] | 28 | Rh | Rh | 1140
| EI | [6 empty] | 6 cm 80% gland | 31 | Mo | Mo | 1180
| Measured EI within +/- 100 of target | [7 empty] | 29 | Mo | Rh | 1150
[9 empty] | 29 | Rh | Rh | 1130
[8 empty] | 8 cm 85% gland | 31 | Mo | Mo | 1220
[9 empty] | 31 | Mo | Rh | 1210
[9 empty] | 31 | Rh | Rh | 1220
Version 2.54
Note on C5: Aili K Bloomquist: If you use the kits sheets you actually have 2.6 mm al, and the test may well fail. Use 4 cm 50-50 instead!
Note on D18: Aili Bloomquist: measure mR through 50-50 or Al.
Note on D19: Aili K Bloomquist: Be sure to change the beam quality offset to match the imaging conditions (see table on right)

## Fuji CR Sensitivity 18
2005-01-01
18 CR Reader Sensitivity - Fuji CR Only
Use technique which gives entrance exposure closest to 20 mR with Mo/Mo
| Patient Name: |  |  |  | Patient ID:
| Target | Mo | [5 empty] | Select the Target S# that matches your mR
| Filter | Mo
| kV Setting | 25 | [5 empty] | mR | S# | +/-
| mAs Setting |  |  |  | Target S-Number | 120 |  | 20 | 120 | 12
| Focal Spot Size | Large |  |  | Target mR | 20 |  | 30 | 80 | 8
| Image Receptor Size |  |  |  | Corrected Target S-Number |  |  | 40 | 60 | 6
Cassette ID
Remove compression paddle.
Make 3 exposures with ion chamber on breast support plate in middle of x-ray field
|  | 1 | 2 | 3 | Average |  |  | Exposure tolerance: | 10%
| mR | [6 empty] | Exposures should be within 10% of each other
| Place freshly erased cassette on breast support plate (above tabletop and grid) and expose | [6 empty] | Use Primary Erasure before exposing image
| Wait 10 minutes and then process under "Sensitivity Test" | [6 empty] | Wait 10 minutes before processing
Before processing/reading plate:
| S Number | [6 empty] | Process under "Sensitivity Test"
If test fails, check that cassette used passed interplate sensitivity test
| S Number should be within 15% 0f Corrected Target S-Number | [6 empty] | Patient name "Physics, Sensitivity"
[8 empty] | S Number Limit | 15%
Version 2.54

## Konica CR Sensitivity 18
2005-01-01
18 CR Reader Sensitivity - Konica CR Only
Use technique which gives entrance exposure closest to 20 mR with Mo/Mo
| Patient Name: |  |  |  |  | Patient ID: |  | Relative Exposure (S)=
| Target | [6 empty] | 2400 | /mR
Filter
kV Setting
mAs Setting
Focal Spot Size
| Image Receptor Size |  |  |  |  | Target Relative Exposure
Cassette ID
Remove compression paddle.
Make 4 exposures with ion chamber on breast support plate in middle of x-ray field, and cassette in holder. Process cassette each time.
Process the cassette under "S Check" each time, and record the Relative Exposure (S value).
|  | 1 | 2 | 3 | 4 | Average |  | Exposure tolerance: | 10%
| mR | [6 empty] | Exposures should be within 10% of each other
Relative Exposure
Use Primary Erasure before exposing image
Wait 10 minutes before processing
Before processing/reading plate:
Process under "S Check"
If test fails, check that cassette used passed interplate sensitivity test
Patient name "Physics, Sensitivity"
[8 empty] | Relative Exposure Limit | 20%
Version 2.54

## Thickness Tracking 2
2005-01-01
Chart 2 - AEC Evaluation | [8 empty] | This can be measured on the
| Phantom |  | 2 cm equiv | 4 cm equiv. | 6 cm equiv. | 8 cm equiv. |  |  |  | fill in while at review station/using software to analyze images
| Type of Attenuator (PMMA or Aluminum) |  | Al | Al | Al | Al |  |  |  | Filled in by the software
| Thickness of Attenuator (mm) |  | 1.27 | 2.54 | 3.8 | 5.1
| AEC mode | [8 empty] | cm equivalent material
| Target | [7 empty] | Tentative SDNR P/F | 2 | 4 | 6 | 8
Filter
| Focal Spot Size |  | Large | Large | Large | Large |  |  | GE 2000D | >3.4, <8.1 | >2.3, <6.9 | >1.8, <5.4 | >1.7, <4.0
| kV Setting | [7 empty] | GE DS | >1.9, <6.7 | >1.6, <4.8 | >1.4, <3.6 | >0.8, <3.3
| mAs | [7 empty] | GE Essential | >5.7, <6.9 | >2.2, <6.0 | >2.3, <4.5 | >2.3, <3.3
| mR | [7 empty] | Siemens Novation | >1.9, < 3.3 | >1.7, <2.9 | >1.5, <2.7 | >0.8, <2.4
| A ROI Contrast Disc | [7 empty] | Siemens Novation Manuf. Spec | >2.4 < 4.2 |  | >1.9 < 2.7
| B ROI Background | [7 empty] | GE SNR Should be greater than 50 for 2 to 6 cm
| C St Dev Background | [8 empty] | P | 1
| SNR | [7 empty] | overall pass/fail summary | F | 2
| SDNR | [8 empty] | M | 3
| Pass/Fail |  | 4 | 4 | 4 | 4 |  | Enter "P" or "F" for pass/fail |  |  | 4
Version 2.54
Density Step Evaluation
| Attenuator: | Al | [6 empty] | To convert Kodak CR Image measurements from original (Log exp.) space to linear space
| AEC Mode: | [7 empty] | ADU_linear = 65535*10^(ADU/1000 -4)
| Step | mAs | [6 empty] | Std_linear = Std*151*10^(ADU/1000-4)
-5
| -3 | [7 empty] | To convert Fuji CR Image measurements from original (Log exp) to linear space
| -1 | [7 empty] | ADU_linear = (10^(L*(ADU-511)/1024)/(S*0.000724) |  |  | This formula converts to "mR" for a 26 kV Mo/Mo beam, assuming 4 cm average breast.
| 0 | [7 empty] | L is latitude (can assume L=2)
| 1 | [7 empty] | S is senstivity number
| 3 | [7 empty] | The transform for Std. Dev is left as an exercise for the reader….;-) (one might be better off just transforming the whole image and then making measurements).
5
| Pass/Fail | 4
Version 2.54
Note on B15: Aili K. Bloomquist: This row is optional - hide it if not using.
Note on D16: Gord Mawdsley: This can usually be measured on acquisition workstation

## Artifact Evaluation 3
Chart 3 - Artifact Evaluation/Flat Field Uniformity - Uniform Phantom | [13 empty] | fill in while at acquisition station/acquiring images
[14 empty] | fill in while at review station/using software to analyze images |  | For CR Use Primary Erasure before exposing image
| Type of Attenuator | PMMA |  |  |  |  | All images acquired "raw" | [6 empty] | Filled in by the software |  | Before processing/reading plate:
| Target Window Width | [14 empty] | Name patient "Physics, Artifacts"
| All Target and Filter combinations should be evaluated. All available focal spot sizes and magnification modes should be evaluated. Evaluate with window set to match that used for Accred. phantom and level set appropriately. | [11 empty] | open new exam for patient "Physics, Artifacts"
| Thickness of Attenuator (mm) | 40 |  |  |  |  | 1.27 | 1.27 | 1.27 | 3.8 |  |  | If detector can operate in two size modes, use the larger one. For CR, check both buckies on most common mode!
| Target | [11 empty] | don't forget 2 and 6 cm images (or 4 and 6 cm if site's phantom not 4 cm thick)
| Filter | [11 empty] | at approprate T/F for acceptance
| Focal Spot Size | Large | Large | Large | Large | Large | Small | Small | Large | Large |  |  | look at raw image, except for FFDM Phantom image, which is viewed with both settings
kV Setting
mAs/mA Setting
| Magnification Factor | 1 | Processed | 1 | 1 | 1 |  |  | 1 | 1
| A - ADU - Disk/Well | [11 empty] | Y | P
| B - ADU - Background | [11 empty] | N | F
| C - Std. Dev. Background | [12 empty] | M
SDNR
Window Width
Window Level
| Artifact Visible? | 2 | 2 | 2 | 3 | 3 | 2 | 3 | 3 | 3 |  |  | Enter "Y" or "N"
| Flat Field Non-Uniformity? | 2 | 2 | 2 | 3 | 3 | 2 | 3 | 3 | 3
| Equipment Artifact? | 2 | 2 | 2 | 3 | 3 | 2 | 3 | 3 | 3
| Detector? | 2 | 2 | 2 | 3 | 3 | 2 | 3 | 3 | 3
| Grid? | 2 | 2 | 2 | 3 | 3 | 2 | 3 | 3 | 3
| Phantom Defect? | 2 | 2 | 2 | 3 | 3 | 2 | 3 | 3 | 3
| Other | 2 | 2 | 2 | 3 | 3 | 2 | 3 | 3 | 3
| Pass/Fail | 1 | 1 | 1 | 4 | 4 | 1 | 4 | 4 | 4 |  | Pass |  |  | close exam
Description of Artifacts:
If significant artifacts are visible, contact the appropriate person maintaining or servicing the digital
| mammography equipment. | [9 empty] | Version 2.54
Use the thickness tracking images, and then image the sites QC phantom at most common AEC setting and also at any clinically relevant target/filter combinations not covered by the thickness tracking test. Also image 2 cm equivalent phantom on mag-stand at site's common mag factor, using clinically relevant target/filter options.
Look at proc. Ver. of first im.

## Dose 6
2005-01-01
Chart 6 - Breast Entrance Exposure, Mean Glandular Dose
fill in while at acquisition station/acquiring images
fill in while at review station/using software to analyze images
|  | Dosimetry system used: |  |  |  |  | Filled in by the software
|  | SID (mm) | 650
|  | Temperature | 1
Barometric Pressure
Use 4.0 cm PMMA (lucite) Flat Field Phantom
#VALUE!
| Mode |  | AEC Auto
kV
| mAs or mA |  |  |  |  | For pre-exposure AEC, the added dose is not taken into account
Target
Filter
| Measured HVL (mm Al) |  | #VALUE! |  |  | NOTE: If the HVL or mR/mAs values don't appear correctly, press CTRL+ALT+F9 to Refresh.
mR/mAs or mR/mA
Calculated Entrance Exposure (mR)
| Conversion Factor mRad/R |  | #VALUE! |  |  | Look up conversion factor in ACR Manual.
| Computed Average Glandular Dose (mGy) |  | #VALUE! |  |  | Displayed dose difference limit: |  | 0.2 | Max Dose: | 3
| Displayed Average Glandular Dose (mGy) |  |  |  |  | Organ dose is dicom tag 0x0040,0x0316, given in dGy.
[5 empty] | Version 2.54 | To convert from dGy to mGy, multiply by 100.
close study
Note on B13: Gord Mawdsley: Standard and Higher Dose mode
Note on D21: Aili K Bloomquist: 4 cm PMMA is deemed equivalent to 4.5 cm 50-50.
Attention Sunnybrook Physicists:
Please enter the following data into the OBSP database (from the main menu, click "Site Information and Visits", find your site in the dropdown menu, click the "Visits" tab, and click the "Equipment Specific" button next to the visit for which you're writing the report):
- Put D20 (Calculated Entrance Exposure) in the "CAR (mR)" field.
- Put D18 (HVL) in the "Measured HVL" field.
- Put D22 (Computed Average Glandular Dose) in the "Dose (mGy)" field.
Delete this box after you've entered the data.

## Ghosting Evaluation 7
2005-01-01
Chart 7 - Evaluation of Ghosting
Suggested patient name "Physics, Ghosting"
Suggested ID number 9905UYYMMDD
Target
Filter
kV Setting
mAs or mA Setting
| Image Receptor Size | [6 empty] | fill in while at acquisition station/acquiring images
fill in while at review station/using software to analyze images
Filled in by the software
A - Mean ROI 1
B - Mean ROI 2
C - Std. ROI 2
Ghost Image SDNR
wait 1 minute between image 1 and image 2
Window Width
Window Level
| Ghost not visible? | 4 | Perform a Secondary Erase on CR plate
Pass
Fail
Marginal
Pass/Fail
Comments
Version 2.54
FFDM Phantom
Setup for Image 1
Tabletop
FFDM Phantom
Setup for Image 2
Delay 60 seconds
Former location of edge of Uniform Phantom
Tabletop
ROI 1, A
ROI 2,
B, C

## Resolution 9
2005-01-01
Chart 9 - MTF - FFDM Phantom | [8 empty] | Acceptable diff from acceptance: |  |  | 10%
|  | Values above reference levels | 4 |  |  |  |  | Pass | suggested patient name "Physics, Phantom"
|  | Values within 10% of acceptance levels | 4 |  |  |  |  | Fail | Suggested ID number 9902UYYMMDD
NA
Use Image from FFDM phantom Test 2.
| Ensure that the phantom image contains no zero or saturated detector values. | [9 empty] | fill in while at acquisition station/acquiring images
fill in while at review station/using software to analyze images
| Use "For Processing" images for calculations if possible. | [9 empty] | Filled in by the software
| Anode | [7 empty] | Use FFDM Accreditation phantom.
| Filter | [7 empty] | Use most commonly used mag-stand position for small focal spot.
| Focal Spot Size | Large | Large | Small | Small |  |  |  | For mag-stand images, ensure MTF tool is surrounded by at least 1 cm of background on all edges
| Mode | man | man
| Mag Factor | 1 | 1 | [5 empty] | For Fuji CR process using "Sensitivity" in the Test Menu
| kV | [7 empty] | For GE, Artifact or Dose phantom technique with twice the mAs works well.
mAs
| Method A - MTF Determination | [7 empty] | Chest wall edge on the left. |  |  |  | DO ROIs in order in quick QC: Left, Right, Up, Down
| x direction | Spatial Frequency | [6 empty] | North = y- |  | up |  | Tentative MTF P/F Criteria |  |  | Mag Stand
| 50% | [7 empty] | South = y+ |  | down |  |  | 50% | 10% | 50% | 10%
| 10% | [7 empty] | East = x+ |  | right |  | GE 2000D | >2.6 | >6.6 | >3.7 | >8.4
| y direction | Spatial Frequency | [6 empty] | West = x- |  | left |  | GE DS | >3.4 | >8 | >4.1 | >7.6
| 50% | [11 empty] | GE Essential | >2.4 | >7.3 | >5.6 | >10
| 10% | [11 empty] | Siemens Novation | >5.1 | >10 | >3.8 | >8.1
| Method B - Manufacturer's Specified Method for Resolution Evaluation | [7 empty] | FUJI Resolution Criteria |  |  | Kodak Resolution Criteria |  |  | GE Resolution Criteria |  |  | Siemens
| Limiting Resolution | [7 empty] | MTF 2 lp | > 40 % | both dirs. | 8 lp/mm visible in both directions |  |  | MTF 2 lp > 49% |  |  | 7 lp/mm visible on 45degree angle
| Parallel to Chest wall | [7 empty] | MTF 4 lp | > 18 % | both dirs. |  |  |  | MTF 4 lp > 18%
Perpendicular to Chest wall
Comments:
close exam
Version 2.54
[6 empty] | 2005-01-01 |  | NB: If Frequency axis looks wrong, you have to manually re-tell the excel chart where the x-coordinates for each series on that plot are (strange excel bug…..) |  |  | Paste MTF Data Here
[11 empty] | MTF -Large-1x | [7 empty] | MTF -Large-1x | [7 empty] | MTF -Small-x | [7 empty] | MTF -Small-x
[8 empty] | x avg | y avg. | Eff. Freq. | Freq | West Edge | East Edge | North Edge | South Edge | x avg | y avg. | Eff. Freq. | Freq | West Edge | East Edge | North Edge | South Edge | x avg | y avg. | Eff. Freq. | Freq | West Edge | East Edge | North Edge | South Edge | x avg | y avg. | Eff. Freq. | Freq | West Edge | East Edge | North Edge | South Edge
[10 empty] | 0 | 0 | [6 empty] | 0 | 0 | [7 empty] | 0 | [7 empty] | 0
Version 2.54
2005-01-01
Version 2.54
Note on B41: Aili K Bloomquist: If freqeuncy axis looks wrong, you must manually re-tel excel where the x-coordinates for that chart are (excel bug)

## Noise and Linearity 10
2005-01-01
can use 40 mm PMMA or 2 mm Al
Chart 10 - Noise and Linearity - Data | [11 empty] | 0
Suggested Patient Name: "Physics, NL"
| Type of Attenuator |  | PMMA | [8 empty] | Suggested ID number 9910UYYMMDD | [7 empty] | Exposure indicator | Signal Linearity check
| Thickness of Attenuator |  | 40 mm | [15 empty] | NA |  | (ADU-offset)/mAs
| Focal Spot Size |  | Large | [9 empty] | fill in while at acquisition station/acquiring images | [5 empty] | Agfa | SAL | SAL/log(mAs)
| Image Receptor Size | [11 empty] | fill in while at review station/using software to analyze images | [5 empty] | Fuji | S | S*mAs
| Target | [11 empty] | Filled in by the software | [5 empty] | Kodak | EI | Linearized Signal / mAs
|  | Filter | [16 empty] | Konica | S | S*mAs
kV
[12 empty] | P/F Thresholds | [8 empty] | Linearized Signal/mAs
| Use "For Processing" Images | [9 empty] | Signal | 20% | Use raw image data for measurements
[11 empty] | SNR | 30%
Detector Signal Uniformity
| Detector Noise Uniformity |  | #VALUE! | [7 empty] | #VALUE!
Variances
|  |  |  | Loc 1 | Loc 2 | Loc 3 | Loc 4 | Loc 5 | Disk |  | Uniformity P/F | mAs or mA | Loc 1 | Loc 2 | Loc 3 | Loc 4 | Loc 5 | log(mAs)
| Low mAs |  | Signal
Std. Dev.
|  |  | SNR | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE!
| Mid mAs |  | Signal
|  |  | Std. Dev. | [8 empty] | mAs or mA | Signal |  | Linearized Signal (may work better for Kodak)
|  |  | SNR | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE!
| Std. mAs |  | Signal
Std. Dev.
|  |  | SNR | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE!
Linear Regressions
| High mAs |  | Signal | [9 empty] | Linear Regression between mAs and Signal | mAs & Signal | log(mAs) & EI
|  |  | Std. Dev. | [9 empty] | Sigma(xy)
|  |  | SNR | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! |  | Sigma(x)
Sigma(y)
Sigma(x^2)
Action Limit: If the background signal value differs by more than +/- 20% from the middle (#1) location, or if the SNR changes by more than +/- 30% from the middle location then the source of the non-uniformity should be identified and corrected.
m
[13 empty] | b |  |  | IF this box does not calculate out OK, experiment with manually pluggin in values for the "Y intercept" to see what gives decent linearity results (around 30 for Siemens?)
If the current SNR changes by more than +/- 10% from the previous SNR measurements then the source of the temporal SNR non-uniformity should be identified and corrected.
Version 2.54
2005-01-01
Chart 11-B - Noise and Linearity - Analysis
Least Squares fit, Noise Linearity
| N/A | Signal Linearity |  |  |  |  | Location | R2
| mAs | [5 empty] | 1
| (ADU-offset)/mAs | [5 empty] | 2
3
4
5
[7 empty] | Overall P/F | N/A
SDNR Analysis
[7 empty] | mAs | SDNR |  |  |  | Least squares fit to noise, min R2 |  |  | 0.95
[12 empty] | Allowed pixel value deviation: |  |  | 10%
min signal lin
max sigal lin
SDNR/sqrt(mAs)
Version 2.54
Note on H57: Aili K. Bloomquist: Possibly should look at Log(mAs) vs SD for CR? Or SDNR?
Chest Wall
2
3
4
5
1
disk
Final pass/fail is at discretion of physicist! Please remove this box when done analysis.

## Geometric Distortion 11
2005-01-01
Chart 11 - Spatial Linearity and Geometric Distortion | [8 empty] | open new exam for patient "Physics, Distortion"
ID number 9913UYYMMDD
| Screen mesh used: | Circuit board |  |  | Mode: | man |  |  | For CR /scanning systems each time |  |  | For Fuji CR, process using "Sensitivity" under the "Test" menu
| kV setting: |  |  |  | Target:
| mAs or mA: |  |  |  | Filter:
| Detector |  | Large FOV | Small FOV |  |  |  | Pass |  | fill in while at acquisition station/acquiring images
| Nominal Image Width (cm) | [6 empty] | Fail
| Nominal Image Height (cm) | [6 empty] | Marginal
| Calculated Image Receptor Size | [8 empty] | fill in while at review station/using software to analyze images
Nominal Width and Height within 5% of Actual
| Uniform Resolution over image area? |  | 4 | 4 | 4 |  |  |  |  | Filled in by the software
| Pattern not distorted? |  | 4 | 4 | 4 |  |  |  |  | GE 2000D,DS | Lorad Selenia | Siemens Selenia | GE LFOV
| Number of Image Columns | [6 empty] | Obtain from DICOM tag 0x0028, 0x0011 |  |  |  | Receptor dimension tolerance |  | 5%
| Number of Image Rows | [6 empty] | Obtain from DICOM tag 0x0028, 0x0010 |  |  |  | pixel size isotropicity tolerance |  | 5%
| Horizontal pixels between Intersections | [6 empty] | RWS |  | Width | Height
| Vertical pixels between Intersections | [6 empty] | RWS | Large FOV
| Effective (tabletop) Pixel Width (micron) | [7 empty] | Small FOV
Effective (tabletop) Pixel Height (micron)
| Pixel Width and Height within 5% | [7 empty] | Nominal Receptor Sizes
[9 empty] | Manuf. | Model | Small | Large
| Pass/Fail | [7 empty] | Fuji | Profect CS/One | 18x24 | 24x30
[9 empty] | GE | 2000D | 19.1x22.9 | NA
| Image Name: | [8 empty] | DS | 19.2x23.04 | NA
| Description of distortion: | [8 empty] | Essential | 19.1x22.9 | 23.9x30.6
[9 empty] | Giotto |  | 17.1x23.9
[9 empty] | Kodak |  | 18x24 | 24x30
| Action Limit: | If there is any significant resolution non-uniformity or pattern distortion, rotate the | [6 empty] | Lorad | Selenia
|  | image on the monitor 90o. If the resolution non-uniformity or distortion | [6 empty] | Planmed | Nuance |  | 23.5x30
|  | persists, seek service from a qualified service engineer to have the problem | [6 empty] | Siemens | Novation | 19.7x21.5 | 23.3x28.6
|  | corrected. | [6 empty] | Siemens | Inspiration | 18 x 24 | 24 x 30
|  | If pixel width vs pixel height differ by more than 5% | [6 empty] | close the exam
or nominal vs actual image dimensions differ by more than 5%
seek service from a qualified service engineer to have the problem corrected.
Version 2.54
Note on B10: Aili K Bloomquist: See manufacturer's specs or dicom tag 0x0018,0x7026 or 0x0018,0x1149 (rows, columns in mm)
Note on B18: Aili K Bloomquist: Obtain from Dicom tag 0x0028, 0x0010
Note on E19: Gord Mawdsley: Measure using workstation ruler. This is column direction, CW-nipple
Note on E20: Aili K Bloomquist: This is row direction, parallel to chest wall

## HARP
HARP Act Semiannual Mammographic COMPLIANCE CHECKLIST | [7 empty] | (DO NOT PRINT)
Items with *** are not legal requirements, however are recommended for the Ontario Breast Screening Program | [7 empty] | Pass/Fail Reference
General Information | [7 empty] | Spreadsheet Link?
Date of tests: |  | 2005-01-01 | [5 empty] | Linked
Institution: | [7 empty] | Linked
Manufacturer: | [5 empty] | Serial #: |  | Linked
Facility
| cassette storage behind shield (***) | [5 empty] | NA | Unlinked
H8.2 -.7) | Quality Control Program | [5 empty] | Y | Unlinked
| Policy and Procedures posted | [5 empty] | Y | Unlinked
Collimation and Alignment
H18.1), R11.1 |  | Mammography
a) beam limitation
|  | i) less than 2% FFD extended to chest wall |  |  |  |  | Y | Unlinked
see data sheet - light field and alignment film
|  | ii) does not extend beyond any other edge of receptor |  |  |  |  | Y | Unlinked
Automatic Exposure Control Design
R7./ H13.1) | Automatic Exposure Control
| a) Indication of selection | [5 empty] | Y | Unlinked
| b) Backup timer |  | 500 | mAs
|  | i) if less than 50 kV, set at 2000 mAs |  |  |  |  | Y | Unlinked
e) If limits hit
|  | i) visible warning |  |  |  |  | Y | Unlinked
|  | ii) manual reset before re-exposure |  |  |  |  | Y | Unlinked
H11.2) | Automatic exposure control terminates at preset radiation exposure | [5 empty] | Y | Unlinked
see data sheet - images to verify reasonable signal and SDNR
for uniform phantom and for varying thickness and kV.
Reproducibility and mR/mAs
R 19.a /H18.4a) | Reproducibility of Exposure
for time > .1 sec (*** use 26 kV,Large, 100 mAs 4.2 phantom)
| i) coefficient of variation < 0.05 | [5 empty] | Y | Unlinked
| ii) all measurements within 15% of mean |  |  |  | see data sheet - Reproducibility |  | Y | Unlinked
R 19.b /H18.3) | kV accuracy
| i) accurate +-1 kV up to 35 | [6 empty] | Linked
R 19.c/H18.4.b) | timer
| ii) accurate to within 7% | [5 empty] | Y | Unlinked
| iii) reproducible at all settings | [5 empty] | Y | Unlinked
|  | i) variation < 0.05 |  |  |  |  | Y | Unlinked
|  | ii) within 15% of mean |  | see data sheet - timer measurement |  |  | Y | Unlinked
R 19.d / H13.4) | mA linearity |  | within 10%
difference between any two adjacent mAs stations
| x1 - x2 <= .1 (x1 + x2) | [6 empty] | Linked
| verify using mR/mAs for all mA stations |  |  |  | see data sheet - Noise and Linearity
H13.5 | mR/mAs in normal range of use | [6 empty] | Linked
difference/sum must be within 10% at any two "mAs" settings differing by factor of two
see data sheet - mAs linearity
Summary
All exposures values are measured at patient entrance under compressor and are free-in-air
All items in the checklist were checked or verified except:
1. The functioning of the line voltage compensator. Design should meet requirements.
2. The applied kilovoltage. A non-invasive meter was used to measure output kV.
3. The true mA and true time. mR/mAs was used to verify linearity, but not accuracy.
Summary of non-compliant items: | [6 empty] | all items on this page comply with HARP regulations
Version 2.54
Note on I11: Aili K Bloomquist: Change for CR

## Technologist Baselines
[7 empty] | 2005-01-01 | 2005-01-01 | [5 empty] | 2005-01-01 | 2005-01-01
ACR Digital Mammography QC Phantom Baseline Calculation Worksheet | [8 empty] | Daily Wet Process Laser Printer Sensitometry Target Values
[10 empty] | Printer |  |  | Film
| FFDM Unit |  |  |  | Room | Think Pink |  |  |  |  | Step / Box # | Date/Time |  |  |  |  | fill in while at acquisition station/acquiring images
|  | AEC Setting | AEC Auto |  | AEC Sensor Position | [8 empty] | Average |  |  | fill in while at review station/using software to analyze images
|  | Target |  |  | Filter |  |  |  |  | D Max | [6 empty] | Filled in by the software
|  | kV |  |  | mAs |  |  |  |  | DD 1
[10 empty] | Mid Density | [5 empty] | Can use "Missed Tissue" image for QC Phantom Baseline if it is of site's phantom, and aquired at clinically relevant technique.
|  |  | Window Width |  | Window Level |  |  |  |  | DD2 | [5 empty] | Use Target, Filter, kV from QC phantom as starting point for imaging uniform phantom for artifacts evaluation baseline
|  | Softcopy | [7 empty] | Base + Fog
Hardcopy
| Op Levels | [10 empty] | Target Value
| mAs | [8 empty] | Density Diff. (DD1-DD2)
| Disk Signal | [8 empty] | Mid Density
| Background Signal | [8 empty] | D Max
| Background St. Dev. | [8 empty] | Base + Fog
|  |  | Target | Lower Eqn | Lower Limit | Upper Eqn | Upper Limit |  |  |  |  | Window width to be used:
|  | mAs |  | mAs X 0.9 |  | mAs X 1.1 | [5 empty] | Window level to be used: |  |  |  | mAs Limit | 0.1
|  | SDNR |  | SDNR X 0.9 |  | SDNR X 1.1 | [9 empty] | SDNR Limit | 0.1
|  | ADU (B) |  | ADU X 0.9 |  | ADU X 1.1 |  |  |  | Reason for setting new target value: (new installation, service to printer, software upgrade, film product change, other) | [5 empty] | ADU Limit | 0.1
Reason for setting new target values: (new installation, service to detector, service to generator, software upgrade, other)
Target ADU from Weekly QC Phantom Image:
| Target |  |  |  | Filter
| kV |  |  |  | mAs
ADUs
| Window Width |  |  |  | Window Level
[8 empty] | Version 2.54 | [6 empty] | Version 2.54
[14 empty] | 2005-01-01 | 2005-01-01
Daily Dry Process Laser Printer Sensitometry Target Values
[10 empty] | Printer |  |  | Film
Step / Box #
D Max
DD 1
Mid Density
DD2
Base + Fog
Target Value
Density Diff. (DD1-DD2)
Mid Density
D Max
Base + Fog
Window width to be used:
Window level to be used:
Reason for setting new target value: (new installation, service to printer, software upgrade, film product change, other)
Version 2.54
Note on B29: Gord Mawdsley: Why is this here? We aren't resetting really.
Note on L5: Aili K Bloomquist: Add box # info from Tech Manual
Note on L41: Aili K Bloomquist: Add box # info from Tech manual

## EndOfTesting
End of Testing - Do not print spreadsheets beyond this point.

## Acceptance
Ontario Breast Screening Program Physics Consulting Group
M.J. Yaffe, Ph.D. | [5 empty] | 480-6100 ext 5715
G.E. Mawdsley, FCCPM, P. Phys. | [5 empty] | 480-6100 ext 5724
FAX(416) 480-5714
Combined HARP and RED ACT Mammographic COMPLIANCE CHECKLIST
Items with *** are not legal requirements, however are recommended for the Ontario Breast
Screening Program
General Information
Date of tests: |  | 2005-01-01
Institution: |  | 0
Mailing address: |  | Sunnybrook HSC, Rm S-632
2075 Bayview Avenue, Toronto ON M4N 3M5
Contact person: |  | 0 |  | Telephone: | 0
Equipment: Mammography Unit
Manufacturer:
Serial number: | 0
Registered owner: |  | 0
Radiation Protection Officer:
Telephone:
Plans registered - |  | Yes, replacement of existing unit, no increase in workload |  |  |  |  | Y
Comments
Version 2.54
Combined HARP and RED Requirements for mammography
Facility
H3.1) | Room shielded - public areas less than .1 mSv/wk |  |  |  | approval |  | Y
H3.5) | Walls and doors shielded to same standard | [5 empty] | Y
H3.6) | Doors fitted with self closers - | [5 empty] | Y
| if accessible to public, "Warning X-Rays" sign | [5 empty] | Y
H3.7) | Film storage protected | [5 empty] | NA
| cassette storage behind shield (***) | [5 empty] | NA
H7.2 -.7) | Quality Control Programme | [5 empty] | Y
| Policy and Procedures posted | [5 empty] | Y
Equipment Supplied and Installed
R3. | installation instructions provided
| a. complete maintenance instructions on site | [5 empty] | Y
| b. | i) maximum line current |  | 3A |  |  | Y
|  | CSA or Hydro approval (***) |  |  | CSA unit, |  | Y
|  | ii) focal spot size nominal |  | L .3 | S .10
|  |  | method of measurement by manufacturer: |  |  |  | Y
|  |  | slit |  |  |  | Y
see data sheet for effective focal spot size (***)
iii) installation and assembly instructions
|  |  | line voltage specs. | 240 |  | 1 ph | Y
|  |  | line impedance specs. |  |  |  | Y
Labelling
H8.1)/R4.a.i | On Control Panel or at Exposure Switch
|  | A) "Unauthorized use" label |  |  |  |  | Y
|  | B) "Hazardous emission in operation" label |  |  |  |  | Y
4.a.ii) | Radiation warning sign - triangle 2cm min | [5 empty] | Y
R4.a.iii) | Identification label | [5 empty] | Y
A) name of manufacturer
B) model designation
|  | C) serial number |  | 0
D) date of manufacture
E) country of manufacture
R4.b. | X-Ray tube housing
R4.b.i) | x-ray tube insert label | [5 empty] | Y
A) name of manufacturer
B) model designation
|  | C) serial number |  | 0
D) date of insert
E) country of manufacture
R4.b.ii) | minimum beam filtration/housing |  |  | 0.8mm Be | 0 mm Al |  | Y
R4.c | X-Ray tube
R4.c.i) | Focal spot mark on tube housing | [5 empty] | Y
| Mark on outside cover (***) | [5 empty] | Y
R4.c.ii) | Cathode and Anode marked (on tube) | [5 empty] | Y
R4.d | Collimator
| total permanent filtration |  |  | 0.8 mm Be | 0 |  | Y
General
R5.1/H8.2 | Every machine shall have:
| H | a) all control, meters clearly identified |  |  |  |  | Y
| R | a) all labels secure |  |  |  |  | Y
| R,H | b) X-ray tube secure and well aligned |  |  |  |  | Y
see data sheet - alignment film #
| R,H | c) Tube housing stable when locked |  |  |  |  | Y
Version 2.54
Meters and Displays
R5.1/H8.2 | d) Warning lights or other indicators | [5 empty] | Y
|  | i) "Ready for exposure" |  |  |  |  | Y
|  | ii) "Exposure" |  |  |  |  | Y
R5.2.c) | meters on adjustable factors
|  |  | A) kVp |  |  |  | Y
B) mA
C) time
|  |  | D) mAs |  |  |  | Y
|  | iii) marks or labels if not adjustable |  |  |  |  | Y
| H | e) Technique indication before exposure |  |  |  |  | Y
| H | f) kVp indicated before exposure if AEC |  |  |  |  | Y
|  |  | special case: (auto kV mode) |  |  |  | Y
R5.2.e) | Audible signal at end of exposure | [5 empty] | Y
Safety
H8.3.a) | exposure + timer controlled by operator | [5 empty] | Y
R5.2.d) | Exposure switch initiate and terminate | [5 empty] | Y
R6.1.3) | Continuous pressure required for x-rays (deadman switch) | [5 empty] | Y
H9.1) | Exposure switch
| a) cannot be conveniently used outside shielded area | [5 empty] | Y
H8.3.b) | filters
| i) in collimator or tube port | [5 empty] | Y
|  | interlock if aluminum below 35 |  |  | OBSP requirement |  | Y
| ii) intercept entire beam | [5 empty] | Y
iii) HVL acceptable
|  | (minimum .3 mm at 30 kVp) |  |  |  |  | Y
see data sheet - HVL
g) filter interlock switches
|  | if i) permanent inherent less than .5 mm Al |  |  |  |  | Y
|  | if ii) used with additional filtration |  |  |  |  | Y
R 5.(2)a) | line voltage compensator | [5 empty] | Y
R 5.(2)b) | if line out of range
| i) warning |  | or ii) interlock |  |  |  | Y
Timer design
H10.1) | Timer terminates at
___a) Preset time interval _X_b) Preset mAs___c) Preset number of pulses
| except with automatic exposure control | [5 empty] | Y
R6.3) | Timing device
| a) Resets to zero at termination of exposure | [5 empty] | Y
| b) Does not permit radiation at "off" position | [5 empty] | Y
Radiation Leakage
R.19.e /H17.1.b) |  | mammographic beam stop transmission
i) mammographic mode, ii) min FFD, iii) max kVp,mAs
< .1 mR (1.1uSv) per exposure
verify with film 500mAs,30 kVp, phantom
see data sheet - light field and alignment
|  | no density under receptor |  |  |  |  | Y
H11.) | Leakage of Beam limiting device < 100 mR/hr(1.1mSv/hr) at 1m | [5 empty] | Y
Survey meter measurement
H13.1) | Housing leakage at 1m < 100 mR/hr (1.1mSv/hr) |  |  |  | <.06 muSv/500 mAs |  | Y
R19.h) | for mammography 5cm from housing less than 2 mR/hr .6mR/500 mAs | [5 empty] | Y
R19.i) | non-exposure conditions, less than 2 mR/hr at 5cm from housing | [5 empty] | Y
R19.j) | no emission from other components over 2 mR/hr at 5cm | [5 empty] | Y
R6.3) | Collimators provide same degree of shielding as tube housing | [5 empty] | Y
Version 2.54
Collimation and Alignment
H17.1), R11.1 |  | Mammography
a) beam limitation
|  | i) less than 2% FFD extended to chest wall |  |  |  |  | Y
see data sheet - Light field and alignment film
|  | ii) does not extend beyond any other edge of receptor |  |  |  |  | Y
H17.2) | Removable cones and diaphragms labelled clearly
| a) | image receptor size |  |  |  |  | Y
| b) | FFD |  |  |  |  | Y
Automatic Exposure Control Design
R7./ H12.1) | Automatic Exposure Control
| a) Indication of selection | [5 empty] | Y
| b) Backup timer |  | Automatically terminates before exposure hits 30 ms.
|  | i) if less than 50 kVp, set at 2000 mAs |  |  |  |  | Y
e) If limits hit
|  | i) visible warning |  |  |  |  | Y
|  | ii) manual reset before re-exposure |  |  |  |  | Y
H10.2) | Automatic exposure control terminates at preset radiation exposure | [5 empty] | Y
see data sheet - images to verify constant signal level
for uniform phantom and for varying thickness and kVp.
Reproducibility and mR/mAs
R 19.a /H17.4a) | Reproducibility of Exposure
for time > .1 sec (*** use 26 kVp,Large, 50 mAs 4.2 phantom)
| i) coefficient of variation < .05 | [5 empty] | Y
| ii) all measurements within 15% of mean | [5 empty] | Y
see data sheet - Reproducibility
R 19.b /H17.3) | kVp accuracy
| i) accurate +-1 kV up to 35 | [5 empty] | Y
| ii) line voltage comp for 7% line variation | [5 empty] | Y
(not tested)
kVp verified using kv meter
|  | see data sheet - kVp Measurement |  |  |  |  | Y
R 19.c/H17.4.b) | timer
| i) can be set to 5 mAs | [5 empty] | Y
| ii) accurate to within 7% | [5 empty] | Y
| iii) reproducible at all settings | [5 empty] | Y
| i) variation < .05 | [5 empty] | Y
| ii) within 15% of mean | [5 empty] | Y
see data sheet - timer Measurement
R 19.d / H12.4) | mA linearity |  | within 10%
difference between any two adjacent mA stations
| x1 - x2 <= .1 (x1 + x2) | [5 empty] | Y
verify using mR/mAs for all mA stations
see data sheet - mA linearity
H12.5 | mR/mAs in normal range of use
difference/sum must be within 10% at any two "mAs" settings differing by factor of two
see data sheet - mAs linearity
Summary
All exposures values are measured at patient entrance under compressor and include backscatter.
All items in the checklist were checked or verified except:
1. The functioning of the line voltage compensator. Design should meet requirements.
2. The applied kilovoltage. A non-invasive meter was used to measure output kVp.
3. The quantitative leakage from the tube at highest workload. Film darkening checked at
10 cm for a large exposure to find gross leaks.
4. The true mA and true time. mR/mAs was used to verify linearity, but not accuracy.
Summary of non-compliant items and corrections:
Version 2.54
Note on A210: Aili Bloomquist: How do we deal with this for Digital?
IF YOU PERFORM THE FOLLOWING ACCEPTANCE TESTS, PLEASE DELETE THIS YELLOW BOX!!
Thank you.

## Laser Printer Eval. 15 AAPM Pat
[7 empty] | 2005-01-01 | 2005-01-01 | 2005-01-01 | Pass | 1
[10 empty] | Fail | 2 |  | fill in while at review station/using software to analyze images
Chart 15-A - Laser Printer - If Available | [9 empty] | Marginal | 3 |  | Filled in by the software
4
Printer:
| 1. Image Quality - TG18 Test Pattern | [5 empty] | Pass |  |  |  |  | Use the TG18-QC pattern on the review workstation
| All steps discernible? |  | 4 | [8 empty] | Assuming: | Ambient Light | 0 | nits
| 0-5 % Target Discernible? |  | 4 | [9 empty] | View Box | 0 | nits
| 95-100% Target Discernible? |  | 4 | [10 empty] | Evenly spaced JND | [8 empty] | Deviation JNDs/Interval
| Patch | % of full scale | OD | Low Contrast? | Use for "outside" versions of AAPM pattern | [7 empty] | Light Output (nits) |  | Expected Luminance | Expected OD | Normalized Contrast | Normalized Expected Contrast | Pass/Fail | lower limit | upper limit | JND
| Black | 100% | [8 empty] | Black is the darker border of the box within a box to the right of patch 1
| 1 | 97.25%
| 2 | 90.98% | [21 empty] | To Calculate JND index as a function of Luminance
| 3 | 84.71% | [21 empty] | j(L) = A + B*Log10(L) + C*(Log10(L))2 + D*(Log10(L))3 + E*(Log10(L))4 + F*(Log10(L))5 + G*(Log10(L))6 + H*(Log10(L))7 + I*(Log10(L))8
| 4 | 78.43% | [21 empty] | A | 71.4981
| 5 | 72.16% | [21 empty] | B | 94.5931
| 6 | 65.88% | [21 empty] | C | 41.9121
| 7 | 59.61% | [21 empty] | D | 9.8247
| 8 | 53.33% | [21 empty] | E | 0.28175
| 9 | 47.06% | [21 empty] | F | -1.1878
| 10 | 40.78% | [21 empty] | G | -0.18014
| 11 | 34.51% | [21 empty] | H | 0.14711
| 12 | 28.24% | [21 empty] | I | -0.017047
| 13 | 21.96%
| 14 | 15.69%
| 15 | 9.412%
| 16 | 3.137%
| White | 0% | [8 empty] | White is the lighter border of the box within a box to the left of patch 16
Viewbox Brightness (nits)
| Ambient light (nits) | [17 empty] | Pass
| Action Limit: | [12 empty] | avg
| If the TG18 test pattern and low contrast targets are not | [12 empty] | stdev
| discernible, the source of change should be identified and corrected. | [10 empty] | To Calculate Luminance as a function of JND index
| 3. Artifact Evaluation | [10 empty] | log10L(j) = (a + c*Ln(j) + e*(Ln(j))2 + g*(Ln(j))3 + m*(Ln(j))4)/(1 + b*Ln(j) + d*(Ln(j))2 + f*(Ln(j))3 + h*(Ln(j))4 + k*(Ln(j))5)
Print uniform image (TG18 UNL-80) with window-width set to 100 % and window-level set to give an OD between 1.5 and 2.0.
[12 empty] | a | -1.3012
| Window |  |  |  | Description of Artifact: | [6 empty] | b | -0.02584
| Level | [10 empty] | c | 0.080243
| OD | [10 empty] | d | -0.1032
| Artifacts acceptable? |  | 4 | [7 empty] | Y | e | 0.13647
| Equipment Artifact? |  | 3 | [7 empty] | N | f | 0.028746
| Other | [10 empty] | g | -0.025468
| If significant artifacts are visible, contact the appropriate person maintaining or servicing the printer | [10 empty] | h | -0.0031979
| or processor. | [7 empty] | Version 2.54 |  |  | k | 0.00012993
[12 empty] | m | 0.0013635
2005-01-01
Chart 15-B - Laser Printer - If Available
|  | Luminance Ratio |  |  |  | L'max/L'min > 100 |  |  |  |  | Luminance Ratio Limit |  | 100
|  | Response |  |  |  | within 20% of GSDF |  |  |  |  | GSDF Response Limit |  | 20%
|  | Jp (ideal JNDs/interval) |  |  |  | < 3 |  |  |  |  | JND Limit |  | 3
|  | Max. Dev. from Jp |  |  |  | < 2 |  |  |  |  | Max Dev |  | 2
|  | RMS Dev. from Jp |  |  |  | < 1 |  |  |  |  | RMS |  | 1
Version 2.54
[7 empty] | 2005-01-01 | 2005-01-01 | 2005-01-01
Chart 15-B - Laser Printer - If Available
| Half Nyquist Frequency Bar Patterns |  |  |  | Nyquist Frequency Bar Patterns
| Location | Direction and Contrast | Visible |  | Location | Direction and Contrast | Visible
| Top Left | H-High | 4 |  | Top Left | H-High | 4
|  | H-Low | 4 |  |  | H-Low | 4
|  | V-High | 4 |  |  | V-High | 4
|  | V-Low | 4 |  |  | V-Low | 4
| Top Right | H-High | 4 |  | Top Right | H-High | 4
|  | H-Low | 4 |  |  | H-Low | 4
|  | V-High | 4 |  |  | V-High | 4
|  | V-Low | 4 |  |  | V-Low | 4
| Middle | H-High | 4 |  | Middle | H-High | 4
|  | H-Low | 4 |  |  | H-Low | 4
|  | V-High | 4 |  |  | V-High | 4
|  | V-Low | 4 |  |  | V-Low | 4
| Bottom left | H-High | 4 |  | Bottom left | H-High | 4
|  | H-Low | 4 |  |  | H-Low | 4
|  | V-High | 4 |  |  | V-High | 4
|  | V-Low | 4 |  |  | V-Low | 4
| Bottom Right | H-High | 4 |  | Bottom Right | H-High | 4
|  | H-Low | 4 |  |  | H-Low | 4
|  | V-High | 4 |  |  | V-High | 4
|  | V-Low | 4 |  |  | V-Low | 4
| Overall Pass/Fail |  |  |  | Overall Pass/Fail
Use "Y" for visible, "N" for not visible.
Version 2.54

## Laser Printer Eval. 15 SMPTE
[6 empty] | 2005-01-01 | 2005-01-01 | 2005-01-01 | Pass | 1
[9 empty] | Fail | 2 |  | fill in while at review station/using software to analyze images
Chart 15-A - Laser Printer - If Available | [8 empty] | Marginal | 3 |  | Filled in by the software
4
Printer:
| 1. Image Quality - SMPTE Test Pattern |  |  |  |  | Pass |  |  |  |  | Use the TG18-QC pattern on the review workstation
| All steps discernible? |  | 4 | [7 empty] | Assuming: | Ambient Light | 0 | nits
| 0-5 % Target Discernible? |  | 4 | [8 empty] | View Box | 0 | nits
| 95-100% Target Discernible? |  | 4 | [9 empty] | Evenly spaced JND | [8 empty] | Deviation JNDs/Interval
| Patch | % of full scale | OD | Use for SMPTE Pattern | [7 empty] | Light Output (nits) |  | Expected Luminance | Expected OD | Normalized Contrast | Normalized Expected Contrast | Pass/Fail | lower limit | upper limit | JND
| Black | 0% | [7 empty] | Black is the darker border of the box within a box to the right of patch 1
| 1 | 10%
| 2 | 20% | [20 empty] | To Calculate JND index as a function of Luminance
| 3 | 30% | [20 empty] | j(L) = A + B*Log10(L) + C*(Log10(L))2 + D*(Log10(L))3 + E*(Log10(L))4 + F*(Log10(L))5 + G*(Log10(L))6 + H*(Log10(L))7 + I*(Log10(L))8
| 4 | 40% | [20 empty] | A | 71.4981
| 5 | 50% | [20 empty] | B | 94.5931
| 6 | 60% | [20 empty] | C | 41.9121
| 7 | 70% | [20 empty] | D | 9.8247
| 8 | 80% | [20 empty] | E | 0.28175
| 9 | 90% | [20 empty] | F | -1.1878
| White | 100% | [7 empty] | White is the lighter border of the box within a box to the left of patch 16 | [12 empty] | G | -0.18014
| Viewbox Brightness (nits) | [21 empty] | H | 0.14711
| Ambient light (nits) | [16 empty] | Pass |  |  |  |  | I | -0.017047
| Action Limit: | [11 empty] | avg
| If the TG18 test pattern and low contrast targets are not | [11 empty] | stdev
| discernible, the source of change should be identified and corrected. | [9 empty] | To Calculate Luminance as a function of JND index
| 3. Artifact Evaluation | [9 empty] | log10L(j) = (a + c*Ln(j) + e*(Ln(j))2 + g*(Ln(j))3 + m*(Ln(j))4)/(1 + b*Ln(j) + d*(Ln(j))2 + f*(Ln(j))3 + h*(Ln(j))4 + k*(Ln(j))5)
Print uniform image (TG18 UNL-80) with window-width set to 100 % and window-level set to give an OD between 1.5 and 2.0.
[11 empty] | a | -1.3012
| Window |  |  | Description of Artifact: | [6 empty] | b | -0.02584
| Level | [9 empty] | c | 0.080243
| OD | [8 empty] | Y | d | -0.1032
| Artifact Visible? |  | 4 | [6 empty] | N | e | 0.13647
| Equipment Artifact? | [9 empty] | f | 0.028746
| Other | [9 empty] | g | -0.025468
| If significant artifacts are visible, contact the appropriate person maintaining or servicing the printer | [9 empty] | h | -0.0031979
| or processor. | [6 empty] | Version 2.54 |  |  | k | 0.00012993
[6 empty] | 2005-01-01 |  |  |  |  | m | 0.0013635
Chart 15-B - Laser Printer - If Available
|  | Luminance Ratio |  |  | L'max/L'min > 100 |  |  |  |  | Luminance Ratio Limit |  | 100
|  | Response |  |  | within 20% of GSDF |  |  |  |  | GSDF Response Limit |  | 20%
|  | Jp (ideal JNDs/interval) |  |  | < 3 |  |  |  |  | JND Limit |  | 3
|  | Max. Dev. from Jp |  |  | < 2 |  |  |  |  | Max Dev |  | 2
|  | RMS Dev. from Jp |  |  | < 1 |  |  |  |  | RMS |  | 1
Version 2.54
Version 2.54

## Laser Printer Eval. 15 Our Pat
[7 empty] | 2005-01-01 | 2005-01-01 | 2005-01-01 | Pass | 1
[10 empty] | Fail | 2 |  | fill in while at review station/using software to analyze images
Chart 15-A - Laser Printer - If Available | [9 empty] | Marginal | 3 |  | Filled in by the software
4
Printer:
| 1. Image Quality - TG18 Test Pattern | [5 empty] | Pass |  |  |  |  | Use the TG18-QC pattern on the review workstation
| All steps discernible? |  | 4 | [8 empty] | Assuming: | Ambient Light | 0 | nits
| 0-5 % Target Discernible? |  | 4 | [9 empty] | View Box | 0 | nits
| 95-100% Target Discernible? |  | 4 | [10 empty] | Evenly spaced JND | [8 empty] | Deviation JNDs/Interval
| Patch | % of full scale | OD | Low Contrast? | Use for Our Version of TG18 Pattern | [7 empty] | Light Output (nits) |  | Expected Luminance | Expected OD | Normalized Contrast | Normalized Expected Contrast | Pass/Fail | lower limit | upper limit | JND
| 1 | 100% | [8 empty] | Black is the darker border of the box within a box to the right of patch 1
| 2 | 94.12%
| 3 | 88.24% | [21 empty] | To Calculate JND index as a function of Luminance
| 4 | 82.35% | [21 empty] | j(L) = A + B*Log10(L) + C*(Log10(L))2 + D*(Log10(L))3 + E*(Log10(L))4 + F*(Log10(L))5 + G*(Log10(L))6 + H*(Log10(L))7 + I*(Log10(L))8
| 5 | 76.47% | [21 empty] | A | 71.4981
| 6 | 70.59% | [21 empty] | B | 94.5931
| 7 | 64.71% | [21 empty] | C | 41.9121
| 8 | 58.82% | [21 empty] | D | 9.8247
| 9 | 52.94% | [21 empty] | E | 0.28175
| 10 | 47.06% | [21 empty] | F | -1.1878
| 11 | 41.18% | [21 empty] | G | -0.18014
| 12 | 35.29% | [21 empty] | H | 0.14711
| 13 | 29.41% | [21 empty] | I | -0.017047
| 14 | 23.53%
| 15 | 17.65%
| 16 | 11.76%
| 17 | 5.882%
| 18 | 0% | [8 empty] | White is the lighter border of the box within a box to the left of patch 16
Viewbox Brightness (nits)
| Ambient light (nits) | [17 empty] | Pass
| Action Limit: | [12 empty] | avg
| If the TG18 test pattern and low contrast targets are not | [12 empty] | stdev
| discernible, the source of change should be identified and corrected. | [10 empty] | To Calculate Luminance as a function of JND index
| 3. Artifact Evaluation | [10 empty] | log10L(j) = (a + c*Ln(j) + e*(Ln(j))2 + g*(Ln(j))3 + m*(Ln(j))4)/(1 + b*Ln(j) + d*(Ln(j))2 + f*(Ln(j))3 + h*(Ln(j))4 + k*(Ln(j))5)
Print uniform image (TG18 UNL-80) with window-width set to 100 % and window-level set to give an OD between 1.5 and 2.0.
[12 empty] | a | -1.3012
| Window |  |  |  | Description of Artifact: | [6 empty] | b | -0.02584
| Level | [10 empty] | c | 0.080243
| OD | [9 empty] | Y | d | -0.1032
| Artifact Visible? |  | 4 | [7 empty] | N | e | 0.13647
| Equipment Artifact? | [10 empty] | f | 0.028746
| Other | [10 empty] | g | -0.025468
| If significant artifacts are visible, contact the appropriate person maintaining or servicing the printer | [10 empty] | h | -0.0031979
| or processor. | [7 empty] | Version 2.54 |  |  | k | 0.00012993
[7 empty] | 2005-01-01 |  |  |  |  | m | 0.0013635
Chart 15-B - Laser Printer - If Available
|  | Luminance Ratio |  |  |  | L'max/L'min > 100 |  |  |  |  | Luminance Ratio Limit |  | 100
|  | Response |  |  |  | within 20% of GSDF |  |  |  |  | GSDF Response Limit |  | 20%
|  | Jp (ideal JNDs/interval) |  |  |  | < 3 |  |  |  |  | JND Limit |  | 3
|  | Max. Dev. from Jp |  |  |  | < 2 |  |  |  |  | Max Dev |  | 2
|  | RMS Dev. from Jp |  |  |  | < 1 |  |  |  |  | RMS |  | 1
Version 2.54
[7 empty] | 2005-01-01 | 2005-01-01 | 2005-01-01
Chart 15-B - Laser Printer - If Available
| Half Nyquist Frequency Bar Patterns |  |  |  | Nyquist Frequency Bar Patterns
| Location | Direction and Contrast | Visible |  | Location | Direction and Contrast | Visible
| Top Left | H-High | 4 |  | Top Left | H-High | 4
|  | H-Low | 4 |  |  | H-Low | 4
|  | V-High | 4 |  |  | V-High | 4
|  | V-Low | 4 |  |  | V-Low | 4
| Top Right | H-High | 4 |  | Top Right | H-High | 4
|  | H-Low | 4 |  |  | H-Low | 4
|  | V-High | 4 |  |  | V-High | 4
|  | V-Low | 4 |  |  | V-Low | 4
| Middle | H-High | 4 |  | Middle | H-High | 4
|  | H-Low | 4 |  |  | H-Low | 4
|  | V-High | 4 |  |  | V-High | 4
|  | V-Low | 4 |  |  | V-Low | 4
| Bottom left | H-High | 4 |  | Bottom left | H-High | 4
|  | H-Low | 4 |  |  | H-Low | 4
|  | V-High | 4 |  |  | V-High | 4
|  | V-Low | 4 |  |  | V-Low | 4
| Bottom Right | H-High | 4 |  | Bottom Right | H-High | 4
|  | H-Low | 4 |  |  | H-Low | 4
|  | V-High | 4 |  |  | V-High | 4
|  | V-Low | 4 |  |  | V-Low | 4
| Overall Pass/Fail |  |  |  | Overall Pass/Fail
Use "Y" for visible, "N" for not visible.
Version 2.54