## Instructions
Notes on the Tech QC Forms version 2.4
All worksheets have been locked against accidentally changing formulas and formatting.
Cells coloured | have been unlocked for user entry of data.
Cells coloured | contain formulas which will automatically fill with values as the sheets are completed.
Graphs have been unlocked to allow the vertical axis scale to be changed which generally improves the graph.
The Technologist QC manual for digital systems contains sufficient instructions to complete these charts.
In the event a critical edit is needed, the password to unprotect any sheet is "obsp".
It is requested that any such change be reported to the OBSP physics consulting group (obsp.physics@sunnybrook.ca)
so that the change may be considered for the master copy.

## Cover
Monthly QC Report Cover Sheet
Facility Name:
Address:
Phone:
Email address:
Excel File Submission
Prepared by:
Date:
Attention: | OBSP Physics Consulting Group
Email Address: | obsp.physics@sunnybrook.ca
Number of Files:
Re: |  | Monthly QC for
Equipment Downtime/Service Report
Other
Message:
Please contact the HEAD QC TECHNOLOGIST as soon as possible if you do not receive all pages.
Name:
Phone: |  |  | ext.
Department:

## Technologist QC List
TABLE 1
Full Field Digital Mammography
Technologist's Quality Control : Minimum Test Frequencies without Laser Printer
Test |  | Minimum Frequency | Corrective Action Timeframe | Relevant Forms
1 | Monitor Cleaning and Daily Checklist | Daily | Immediately, before checked component is used for patients | Chart A, Chart 1
3 | Phantom Image Quality | Weekly | Immediately | Chart 4
4 | Display Monitor QC | Weekly | Immediately: workstation-before patient images interpreted; acquisition station monitor- before patients imaged | Chart 5, Chart 6
5 | Viewbox Cleanliness | Weekly | Immediately, before patient images interpreted or comparison films reviewed | Chart A
6 | Full Field Artefacts* | Monthly | Immediately | Chart 7
7 | Monthly Checklist | Monthly | Immediately or within 30 days, depending on check | Chart 8
9 | Resolution/MTF~ | Monthly | Immediately | Chart B
10 | Repeat Analysis | Quarterly | Within 30 days of the test date | Chart 10, Chart 11
13 | Compression Force Test | Semi-Annually | Immediately | Chart B
*All CR systems
~All systems with moving parts in the detector
TABLE 2
Full Field Digital Mammography
Technologist's Quality Control : Minimum Test Frequencies with Laser Printer
Test |  | Minimum Frequency | Corrective Action Timeframe | Relevant Forms
1 | Monitor Cleaning and Daily Checklist | Daily | Immediately, before checked component is used for patients | Chart A, Chart 1
2 | Laser Printer Sensitometry+^ | Daily+ | Immediately, before patient images printed | Chart 2, Chart 3
Monthly^
3 | Phantom Image Quality | Weekly | Immediately | Chart 4
4 | Display Monitor QC | Weekly | Immediately: workstation-before patient images interpreted; acquisition station monitor- before patients imaged | Chart 5, Chart 6
5 | Viewbox Cleanliness | Weekly | Immediately, before patient images interpreted or comparison films reviewed | Chart A
6 | Full Field Artefacts* | Monthly | Immediately | Chart 7
7 | Monthly Checklist | Monthly | Immediately or within 30 days, depending on check | Chart 8
8 | Laser Printer Artefacts | Monthly | Immediately, before patient images printed | Chart 9
9 | Resolution/MTF~ | Monthly | Immediately | Chart B
10 | Repeat Analysis | Quarterly | Within 30 days of the test date | Chart 10, Chart 11
11 | Printed Image Quality Test | Quarterly | Immediately, before patient images printed | Chart 12
12 | Analysis of Fixer Retention+ | Quarterly | Within 30 days of the test date | Chart B
13 | Compression Force Test | Semi-Annually | Immediately | Chart B
+All systems with wet processing
^All systems with dry processing
*All CR systems
~All systems with moving parts in the detector
**All printers with wet processing and without daylight loading magazines or integrated processing

## Daily, Wkly Chart A
Chart A: Digital Mammography Quality Control Checklist
Daily and Weekly Tests
|  | Room: | [15 empty] | Month: | [8 empty] | Year:
Checkmark (√) = Pass/Adequate; x = Fail; initial when complete
| Test | Date | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31
Initials
| 1 | Monitor Cleaning
| 1 | Daily Checklist (daily)
| 2 | Laser Printer Sensi-tometry-wet (daily)
| 3 | Phantom Image Quality (weekly)
| 4 | Display Monitor QC (weekly)
| 5 | Viewbox Cleanliness (weekly)
| Date: |  |  | Test: | [8 empty] | Comments:

## Mnthly, Qrtly, SAnn Chart B
Chart B: Digital Mammography Quality Control Checklist
Monthly, Quarterly and Semi-Annual Tests
| Room: |  |  |  | Year: | [5 empty] | Minimum Resolution:
Test | Month | JAN | FEB | MAR | APR | MAY | JUN | JUL | AUG | SEP | OCT | NOV | DEC
Pass (P) / Fail (F)
Initials
2 | Laser Printer Sensitometry - dry (monthly)
6 | Full-Field Artefacts (CR-monthly)
7 | Monthly Checklist (monthly)
8 | Laser Printer Artefacts (monthly)
9 | Resolution/MTF Horizontal Bars (scanning-quarterly)
Resolution/MTF Vertical Bars (scanning-quarterly)
Overall Pass/Fail
10 | Repeat Analysis (quarterly)
11 | Printed Image Quality (quarterly)
12 | Analysis of Fixer Retention (quarterly)
13 | Compression Force (semi-annual)
Date: |  |  | Test: |  |  |  |  | Comments:

## Technique Chart C
Chart C: Full Field Digital Mammography Technique Chart
No Implants
Compressed Breast Thickness (cm) | Mammographic Density | AEC Setting | Target | Filter | kVp | mAs | Aim Value (S#, EI or ADU)
Implants
Compressed Breast Thickness (cm) | Mammographic Density | AEC Setting | Target | Filter | kVp | mAs | Aim Value (S#, EI or ADU)

## Daily Checklist 1
CHART 1: DAILY CHECKLIST
Facility: | [24 empty] | Room:
Month: | [24 empty] | Year:
Checkmark (√) = Pass/Adequate; x = Fail; initial when complete
Day --> | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31
Initials
No Loose Parts
Overall Integrity
Cleanliness
No Cracks in Paddles
Hoses and Cabling Unobstructed
CR-Reader Cleanliness
Imaging Plates Erased
Date: |  | Comments:

## Laser Printer 2
CHART 2: LASER PRINTER SENSITOMETRY
Laser Printer: |  |  | Film: |  |  | Emul#: |  |  | Workstation:
TG-18 QC | Box #18 | Box #16 | Box #11 | Box #5 | Box #16 - Box #5 | Box #1 | WW: | WL:
Gray-Scale Step Wedge | Step #: | Step #: | Step #: | Step #: | DD1-DD2 | Lightest
Op Levels
Date | Dmax | DD1 | MD | DD2 | DD | B+F | Remarks: |  |  | DMaxMid | DMaxLL | DDmid | DDLL | DDUL | MDmid | MDLL | MDUL | BFmid | BFUL

## Laser QC Baselines 3
CHART 3: SETTING OPERATING LEVELS FOR LASER PRINTER
Date operating levels valid from:
TG-18 QC Pattern | Window Width: |  | Window Level:
| Box#/ Step | Dates |  | Average | Operating Level
Maximum Density | 18
DD1 | 16
Mid-Density | 11
DD2 | 5
Base+Fog | 1
DD
Reason for setting new operating levels:

## Weekly Phantom 4
CHART 4: WEEKLY PHANTOM
Facility:
Room: |  |  |  | Manuf: |  |  |  |  | Model:
TG-18 QC | mAs | Disk Signal (A) | Background Signal (B) | Background St. Dev. (C) | SDNR (B-A)/C | Artefacts (Y/N) | EI or S (CR only) | WW: | WL:
Op Levels
Date | mAs | Disk Sig | Backgd. Sig. | Std. Dev. | DD | Artefacts | EI or S | Remarks |  | mAsmid | mAsLL | mAsUL | SIGmid | SIGLL | SIGUL | SDNRmid | SDNRLL | SDNRUL | EImid | EILL | EIUL

## Monitor QC 5
CHART 5: WEEKLY DISPLAY MONITOR QC - REVIEW WORKSTATION
Facility: | [7 empty] | Room:
Workstation Mfr: | [7 empty] | Monitor Mfr:
Year:
Month
Day
Initials
Monitor | Left | Right | Left | Right | Left | Right | Left | Right | Left | Right | Left | Right
General Image Quality (Pass(P)/Fail(F)/Marginal(M))
No Smearing
No Artefacts
Ramps cont.
Geometric Distortion
Lines straight
Pattern Centered
Boxes square
Luminance
Patches distinct
0-5%
95-100%
# of Letters Visible (at least 11 or "QUALITY CONT")
Dark
Mid-Gray
Light
Clinical Image Check
Background (non-breast) area is black
Background areas on 2 monitors match
Dense breast tissues on 2 monitors match
Contrast on 2 monitors matches
Overall Pass/Fail
Remarks
Date: |  | Action:
Under "General Image Quality" enter "Y" or "N" to indicate presence or absence of smearing or other artefacts.
Enter "Y"if the greyscale ramps appear continuous, "N" if they do not.
Under "Geometric Distortion" enter "Y" or "N" to indicate if the lines and borders appear straight,
if the pattern is centered and if the boxes appear square.
Under "Luminance" enter "Y" or "N" to indicate the visibility of each item.
Under "Lettering" record the number of letters visible at each luminance level.

## Monitor QC 6
CHART 6: WEEKLY DISPLAY MONITOR QC - ACQUISITON MONITOR
Facility: |  | Room:
Workstation Mfr: |  | Monitor Mfr:
Year:
Month
Day
Initials
General Image Quality (Pass(P)/Fail(F)/Marginal(M))
No Smearing
No Artefacts
Ramps cont.
Geometric Distortion
Lines straight
Pattern Centered
Boxes square
Luminance
Patches distinct
0-5%
95-100%
Mid-Gray
Light
Overall Pass/Fail
Remarks
Date: | Action:
Under "General Image Quality" enter "Y" or "N" to indicate presence or absence of smearing or other artefacts.
Enter "Y"if the greyscale ramps appear continuous, "N" if they do not.
Under "Geometric Distortion" enter "Y" or "N" to indicate if the lines and borders appear straight, if the pattern is
centered and if the boxes appear square.
Under "Luminance" enter "Y" or "N" to indicate the visibility of each item.
Under "Lettering" record the number of letters visible at each luminance level.

## Artifacts 7
CHART 7 - CR MONTHLY FULL-FIELD ARTEFACT EVALUATION
Facility: |  | Room:
Window Width: |  | Year:
Window Level:
Compression Force |  | AEC Mode:
Month
Day
Initials
Pass (P) / Fail (F) / Marginal (M)
Cassette ID
Target
Filter
kV
mAs
No Undue Artefacts
Hardcopy (H) /Softcopy (S) / Both (B)
Type of copy
Date | Remarks

## Visual Checklist 8
CHART 8: MONTHLY MECHANICAL INSPECTION
Facility: | [7 empty] | Room:
Unit: | [7 empty] | Year:
Month | Jan | Feb | Mar | Apr | May | Jun | Jul | Aug | Sep | Oct | Nov | Dec
Date
Initials
Pass (P) / Fail (F) / Marginal (M) / Not Available (NA)
Room temperature
No Loose Parts
Cleanliness
No Cracks in Paddle
Overall Integrity
Hoses and Cabling Unobstructed
Angulation Indicator
Locks (all)
Field Light
Smoothness of Motion
Panel switches /lights/meters
Time and Date on images correct
Facility ID on images correct
Technique charts
Cleaning solution
IF failure of a test listed above, document resolution:

## Laser Printer Artifacts 9
CHART 9: MONTHLY LASER PRINTER ARTEFACT EVALUATION
Facility: | [7 empty] | Year:
Printer:
Month | Jan | Feb | Mar | Apr | May | Jun | Jul | Aug | Sep | Oct | Nov | Dec
Date
Initials
OD
Pass (P) / Fail (F) / Marginal (M) / Not Available (NA)
Uniform OD
No Streaks?
No Lines?
No Specks?
No Other Artefacts?
Overall Pass/Fail
IF failure of a test listed above, document resolution:

## Retake Record 10
CHART 10: DIGITAL MAMMOGRAPHY RETAKE RECORD
(Enter any repeated exposures that required the patient to have additional dose beyond that of the normal exam)
Facility: |  |  |  | Room:
Unit:
Date From:
Date To:
Study # | Causes | Number of Times | Date | Initiated by: |  | Technologist
|  |  |  | Physician | Technologist
Causes:
1 | Positioning
2 | Patient Motion
3 | Improper Detector Exposure
4 | Artefact
5 | X-Ray Equipment Failure
6 | Software Failure
7 | Blank Image
8 | No Image
9 | Other

## Retake Analysis 11
CHART 11: QUARTERLY RETAKE ANALYSIS
Facility: |  |  | Room:
Date From: |  |  | Unit:
Date To:
Total Patient Exposures in Quarter:
| Cause | No. of Exposures | Percentage of Repeats
1 | Positioning
2 | Patient Motion
3 | Improper Detector Exposure
4 | Artefact
5 | X-Ray Equipment Failure
6 | Software Failure
7 | Blank Image
8 | No Image
9 | Other
|  | Total | % Repeats
Repeat Exposures (1-9)

## Printed Image Quality 12
CHART 12: QUARTERLY PRINTED IMAGE QUALITY TEST
Printer: | Year:
Printed From:
Date
Initials
Length of 5 cm Horizontal Ruler (between 4.7 and 5.3 cm)
Length of 5 cm Vertical Ruler (between 4.7 and 5.3 cm)
Pass (P) / Fail (F) / Marginal (M) / Not Available (NA)
0-5% Square
95-100% Square
Horizontal Line Pairs
Vertical Line Pairs
Grey Patches
Lines Straight
No Artefacts
Overall Pass/Fail
IF failure of a test listed above, document resolution: